## Summary
No. | Model | Serial | Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
1 | EG-530WR | 1G361K705 | Receive | 2015-12-10 11:23 | 2016-11-02 11:20 | 2017-07-28 17:01 | 2019-01-18 11:38 | 2021-01-12 11:42 | 2022-04-04 15:04
|  |  | Repair | 2015-12-10 11:23: 2015-12-11 11:14; 2017-07-28 17:01: 2017-09-15 08:33
|  |  | Size | Minor | Minor | Major | Major | Major | Major
|  |  | Issue | E | RBS | LGB, OSA | CHA, FSB, BSA, FCT | CHA, OSA, FSB, BSA, FCT, FSA, VCA | CHA, OSA, FSB, BSA, FCT, FSA, VCA
|  |  | RMA# | FMSV2015120006 | FMSV2016110003 | FMSV2017070033 | FMSV2019010015 | FMSV2021010019 | FMSV2022040007
|  |  | Defect | The OSA main case was cracked and leaked | 1. RBS is leaked._x000D_ 2. BSA is deformed._x000D_ 3. VCA slide sleeve is damaged. |  | CHA is cracked_x000D_ Leaking at rubber cone of FSB._x000D_ BSA is deformed_x000D_ FCT is restricted | 1. OBL cracked_x000D_ 2. Rubber cone FSA leaked_x000D_ 3. Rubber cone FSB cracked_x000D_ 4. BSA deformed_x000D_ 5. FSA deformed and scratched._x000D_ 6. Belt for VCA Cap broken | 1. OBL cracked_x000D_ 2. Rubber cone FSA leaked_x000D_ 3. Rubber cone FSB cracked_x000D_ 4. BSA deformed_x000D_ 5. FSA deformed and scratched._x000D_ 6. Belt for VCA Cap broken
|  |  | Parts |  | RBS-G166A | BSA-G166A RUBBER FOR BSA ROHS SLEEVE ROHS NUT ASSY ROHS FCT RING FG381A-6 ROHS FORCEPS CHANNEL TUBE AIR TUBE W-TUBE FG246A ROHS LGB-G201A ROHS LG LENS ASS/Y FX 35A ROHS | CHA-G201A ROHS AIR TUBE W-TUBE FG246A ROHS FORCEPS CHANNEL TUBE FCT RING FG381A-6 ROHS LG FLEXIBLE SECTION ASSY ROHS BSA-G166A SLEEVE ROHS NUT ASSY ROHS RUBBER FOR BSA ROHS | ISA-G201A-6(AT/WT) SLEEVE ROHS NUT ASSY LG FLEXIBLE SECTION ASSY ROHS BELT ROHS | ISA-G201A-6(AT/WT) SLEEVE ROHS NUT ASS'Y BELT LG FLEXIBLE SECTION ASSY ROHS
|  |  | Approval | Approval | Decline | Approval | Cancel | Cancel | Cancel
2 | EG-530WR | 1G361K775 | Receive | 2017-02-28 10:23 | 2018-05-28 17:36 | 2018-06-28 16:03 | 2018-10-08 13:29 | 2018-11-13 11:35 | 2019-11-29 14:41 | 2020-12-10 10:36 | 2022-04-04 15:04
|  |  | Repair | 2017-03-13 08:59 | 2018-06-22 09:27 | 2018-07-24 13:41 | 2018-10-11 13:19 | 2019-02-28 09:24 | 2020-01-15 08:27
|  |  | Size | Major | Minor | Minor | Minor |  | Minor | Major | Major
|  |  | Issue | RBS, LGB | AWT | AWT | AWT | VCA, BSA, LPA, FCT | AWT | CHA, BSA, FSA, LGB | CHA, RBS, BSA, FSA, LGB
|  |  | RMA# | FMSV2017020007 | FMSV2018050060 | FMSV2018060035 | FMSV2018100007 | FMSV2018110018 | FMSV2019110070 | FMSV2020120021 | FMSV2022040006
|  |  | Defect | 1. Leaked at RBS_x000D_ 2. LGB is broken 1 side | AWT clogged | AW channel clogged, LG broken 01 side |  | LGB is slipped into the Bending section_x000D_ BSA deformed_x000D_ Can not connect to Processor VP4450 because broken cable of VCA_x000D_ VCA lossen when connect to Processor |  | 1. Stain on image._x000D_ 2. Leaked at RBS._x000D_ 3. BSA deformation._x000D_ 4. FSA worn out._x000D_ 5. LGB uneven illumination. | 1. Stain on image._x000D_ 2. Leaked at RBS._x000D_ 3. BSA deformed_x000D_ 4. FSA scratched_x000D_ 5. LGB broken
|  |  | RMA# | 2017-03-13 08:59: LG BUNDLE RBS-G166A; 2019-02-28 09:24: BSA-G166A RUBBER FOR BSA ROHS SLEEVE ROHS NUT ASSY ROHS FORCEPS CHANNEL TUBE FCT RING FG381A-6 ROHS AIR TUBE W-TUBE FG246A ROHS SLIDE SLEEVE | ISA-G290A SLEEVE ROHS NUT ASSY | ISA-G201A-6(AT/WT) SLEEVE ROHS NUT ASSY
|  |  | Approval | Approval | Approval | Approval | Approval | Approval | Approval | Cancel | Cancel
3 | EG-530WR | 2G361K530 | Receive | 2018-11-20 12:09 | 2019-03-15 13:46 | 2019-07-29 11:37 | 2020-09-03 17:05
|  |  | Repair | 2017-03-13 08:59: 2018-11-20 13:14; 2018-06-22 09:27: 2019-03-29 16:46; 2018-07-24 13:41: 2019-10-03 09:17
|  |  | Size | Minor | Minor | Major | Major
|  |  | RMA# | RPR, O | RBS | CHA, FCT, FSA | RBS, CHA, FCT, BSA, FSA
|  |  | RMA# | FMSV2018110032 | FMSV2019030021 | FMSV2019070052 | FMSV2020090009
|  |  | Defect | 2018-06-22 09:27: Leaking at RBS; 2018-07-24 13:41: - FCT broken_x000D_ - Stain on the image_x000D_ - FSA deformed; 2018-10-11 13:19: 1. Air leakege at RBS._x000D_ 2. Stain & black spots on the image._x000D_ 3. FST restricted_x000D_ 4. BSA & FSA deformed.
|  |  | Parts | 2018-06-22 09:27: RUBBER FOR BSA ROHS; 2018-07-24 13:41: ISA-G201A-6(AT/WT) ISA RING 2 SLEEVE ROHS NUT ASSY ROHS; 2018-10-11 13:19: CAMERA HEAD ASS''Y FCT-G76D RING CAP +UV ROHS NOZZLE AIR TUBE WATER TUBE BSA-G166A SLEEVE ROHS NUT ASSY FSA-G166A ISA RING 2 FSA RING RUBBER FOR BSA ROHS
|  |  | Approval | Approval | Approval | Approval | Decline
4 | EG-530NW | 2G366K004 | RMA# | 2017-03-13 08:59: 2020-10-05 16:49; 2018-06-22 09:27: 2021-03-10 14:07
|  |  | Repair | 2017-03-13 08:59: 2020-10-08 15:56; 2018-06-22 09:27: 2021-03-12 09:13
|  |  | Size | 2017-03-13 08:59: Minor; 2018-06-22 09:27: Minor
|  |  | Issue | 2017-03-13 08:59: AWT, OSA; 2018-06-22 09:27: AWT
|  |  | RMA# | 2017-03-13 08:59: FMSV2020100011; 2018-06-22 09:27: FMSV2021030040
|  |  | RMA# | 2017-03-13 08:59: 1. Clogged AWJ._x000D_ 2. U/D loosen.; 2018-06-22 09:27: 1. AW restricted (priority)_x000D_ 2. LGB broken 01 side_x000D_ 3. bsa defomred .Customer just want AW become normal
Parts
|  |  | Approval | 2017-03-13 08:59: Approval; 2018-06-22 09:27: Approval
5 | EG-600WR | 5G391K103 | Receive | 2017-03-13 08:59: 2019-09-10 11:50; 2018-06-22 09:27: 2020-10-05 16:55; 2018-07-24 13:41: 2021-03-10 14:08
|  |  | Repair | 2017-03-13 08:59: 2019-09-11 09:52; 2018-06-22 09:27: 2020-10-07 15:34; 2018-07-24 13:41: 2021-03-12 08:12
|  |  | RMA# | 2017-03-13 08:59: Minor; 2018-06-22 09:27: Minor; 2018-07-24 13:41: Minor
|  |  | Issue | 2017-03-13 08:59: CHA, AWT; 2018-06-22 09:27: AWT, OSA; 2018-07-24 13:41: AWT
|  |  | RMA# | 2017-03-13 08:59: FMSV2019090019; 2018-06-22 09:27: FMSV2020100012; 2018-07-24 13:41: FMSV2021030041
|  |  | Defect | 2018-06-22 09:27: 1. Clogged AWT and WJ._x000D_ 2. U/D/L/R loosen.; 2018-07-24 13:41: Clogged AWJ, AWT
Parts
|  |  | Approval | 2017-03-13 08:59: Approval; 2018-06-22 09:27: Approval; 2018-07-24 13:41: Approval
6 | EG-530WR | 6G361K178 | Receive | 2017-03-13 08:59: 2020-10-26 15:03
|  |  | Repair | 2017-03-13 08:59: 2020-10-28 13:54
|  |  | Size | 2017-03-13 08:59: Minor
|  |  | Issue | 2017-03-13 08:59: AWT
|  |  | RMA# | 2017-03-13 08:59: FMSV2020100059
Defect
Parts
|  |  | Approval | 2017-03-13 08:59: Approval
7 | EG-530WR | 6G361K179 | Receive | 2017-03-13 08:59: 2019-10-28 09:34; 2018-06-22 09:27: 2020-10-15 15:26; 2018-07-24 13:41: 2022-12-05 16:00
|  |  | Repair | 2017-03-13 08:59: 2019-11-01 10:08; 2018-06-22 09:27: 2020-10-22 16:54
|  |  | Size | 2017-03-13 08:59: Minor; 2018-06-22 09:27: Minor; 2018-07-24 13:41: Major
|  |  | Issue | 2017-03-13 08:59: AWT, ANGL; 2018-06-22 09:27: AWT; 2018-07-24 13:41: CHA, SW, BSA, FSA
|  |  | RMA# | 2017-03-13 08:59: FMSV2019100043; 2018-06-22 09:27: FMSV2020100043; 2018-07-24 13:41: FMSV2022120005
|  |  | Defect | 2018-06-22 09:27: Clogged AWJ; 2018-07-24 13:41: 1. Noise on image_x000D_ 2. Rainbow on image_x000D_ 3. RC switch broken_x000D_ 4. BSA and FSA deformed_x000D_
|  |  | Parts | 2018-07-24 13:41: ISA-G201A-6(AT/WT) SLEEVE ROHS NUT ASS'Y RC SWITCH SW CAP
|  |  | Approval | 2017-03-13 08:59: Approval; 2018-06-22 09:27: Approval; 2018-07-24 13:41: Decline

## Installation
National Hospital of Tropical Diseases
No. | Model | SN | Date Installed | WTY_START | WTY_END
1 | EB-530T | 1B084K360 | 2016-10-20 | 2016-10-20 | 2017-10-20
2 | EG-530WR | 1G361K075 | 2015-08-09 | 2015-08-09 | 2016-08-08
3 | EG-530WR | 1G361K705 | 2015-08-09 | 2015-08-09 | 2016-08-08
4 | EG-530WR | 1G361K775 | 2015-08-09 | 2015-08-09 | 2016-08-08
5 | VP-3500 | 1V609K184 | 2015-08-09 | 2015-08-09 | 2016-08-08
6 | VP-3500 | 1V609K844 | 2016-10-20 | 2016-10-20 | 2017-10-20
7 | VP-3500 | 1V609K845 | 2016-10-20 | 2016-10-20 | 2017-10-20
8 | EG-530WR | 2G361K530 | 2016-10-20 | 2016-10-20 | 2017-10-20
9 | EG-530NW | 2G366K004 | 2019-05-20 | 2019-05-20 | 2020-05-20
10 | XL-4450 | 2S094K059 | 2015-08-09 | 2015-08-09 | 2016-08-08
11 | VP-3500 | 2V609K813 | 2018-12-26 | 2018-12-26 | 2019-12-26
12 | VP-3500 | 2V609K814 | 2018-12-26 | 2018-12-26 | 2019-12-26
13 | EC-530WL3 | 3C643K260 | 2019-05-20 | 2019-05-20 | 2020-05-20
14 | XL-4450 | 3S094K029 | 2016-10-20 | 2016-10-20 | 2017-10-20
15 | XL-4450 | 3S094K147 | 2016-10-20 | 2016-10-20 | 2017-10-20
16 | EB-530P | 4B085K030 | 2019-05-20 | 2019-05-20 | 2020-05-20
17 | XL-4450 | 4S094K337 | 2019-05-20 | 2019-05-20 | 2020-05-20
18 | EG-600WR | 5G391K103 | 2019-05-20 | 2019-05-20 | 2020-05-20
19 | EG-530WR | 6G361K178 | 2019-05-20 | 2019-05-20 | 2020-05-20
20 | EG-530WR | 6G361K179 | 2019-05-20 | 2019-05-20 | 2020-05-20

## Sheet1
RMA No. | SERIAL_NO | MODEL | RECIEVE_DATE | Repair Date | Repair Size | Defect Note | Approval | Repair | No#
FMSV2015120006 | 1G361K705 | EG-530WR | 2015-12-10 11:22:32 | 2015-12-11 11:13:49 | Minor | The OSA main case was cracked and leaked | Approval | E | 1
FMSV2016110003 | 1G361K705 | EG-530WR | 2016-11-02 11:20:21 |  | Minor | 1. RBS is leaked. 2. BSA is deformed. 3. VCA slide sleeve is damaged. | Decline | RBS | 2
FMSV2017070033 | 1G361K705 | EG-530WR | 2017-07-28 17:00:51 | 2017-09-15 08:33:17 | Major |  | Approval | LGB, OSA | 3
FMSV2019010015 | 1G361K705 | EG-530WR | 2019-01-18 11:38:06 |  | Major | CHA is cracked Leaking at rubber cone of FSB. BSA is deformed FCT is restricted | Cancel | CHA, FSB, BSA, FCT | 4
FMSV2021010019 | 1G361K705 | EG-530WR | 2021-01-12 11:42:16 |  | Major | 1. OBL cracked 2. Rubber cone FSA leaked 3. Rubber cone FSB cracked 4. BSA deformed 5. FSA deformed and scratched. 6. Belt for VCA Cap broken | Cancel | CHA, OSA, FSB, BSA, FCT, FSA, VCA | 5
FMSV2022040007 | 1G361K705 | EG-530WR | 2022-04-04 15:04:12 |  | Major | 1. OBL cracked 2. Rubber cone FSA leaked 3. Rubber cone FSB cracked 4. BSA deformed 5. FSA deformed and scratched. 6. Belt for VCA Cap broken | Cancel | CHA, OSA, FSB, BSA, FCT, FSA, VCA | 6
FMSV2017020007 | 1G361K775 | EG-530WR | 2017-02-28 10:22:37 | 2017-03-13 08:58:42 | Major | 1. Leaked at RBS 2. LGB is broken 1 side | Approval | RBS, LGB | 1
FMSV2018050060 | 1G361K775 | EG-530WR | 2018-05-28 17:36:30 | 2018-06-22 09:27:30 | Minor | AWT clogged | Approval | AWT | 2
FMSV2018060035 | 1G361K775 | EG-530WR | 2018-06-28 16:02:49 | 2018-07-24 13:40:51 | Minor | AW channel clogged, LG broken 01 side | Approval | AWT | 3
FMSV2018100007 | 1G361K775 | EG-530WR | 2018-10-08 13:29:28 | 2018-10-11 13:19:06 | Minor |  | Approval | AWT | 4
FMSV2018110018 | 1G361K775 | EG-530WR | 2018-11-13 11:34:35 | 2019-02-28 09:24:04 |  | LGB is slipped into the Bending section BSA deformed Can not connect to Processor VP4450 because broken cable of VCA VCA lossen when connect to Processor | Approval | VCA, BSA, LPA, FCT | 5
FMSV2019110070 | 1G361K775 | EG-530WR | 2019-11-29 14:41:01 | 2020-01-15 08:26:31 | Minor |  | Approval | AWT | 6
FMSV2020120021 | 1G361K775 | EG-530WR | 2020-12-10 10:36:01 |  | Major | 1. Stain on image. 2. Leaked at RBS. 3. BSA deformation. 4. FSA worn out. 5. LGB uneven illumination. | Cancel | CHA, BSA, FSA, LGB | 7
FMSV2022040006 | 1G361K775 | EG-530WR | 2022-04-04 15:03:33 |  | Major | 1. Stain on image. 2. Leaked at RBS. 3. BSA deformed 4. FSA scratched 5. LGB broken | Cancel | CHA, RBS, BSA, FSA, LGB | 8
FMSV2018110032 | 2G361K530 | EG-530WR | 2018-11-20 12:09:08 | 2018-11-20 13:14:07 | Minor |  | Approval | RPR, O | 1
FMSV2019030021 | 2G361K530 | EG-530WR | 2019-03-15 13:46:02 | 2019-03-29 16:46:08 | Minor | Leaking at RBS | Approval | RBS | 2
FMSV2019070052 | 2G361K530 | EG-530WR | 2019-07-29 11:36:57 | 2019-10-03 09:17:18 | Major | - FCT broken - Stain on the image - FSA deformed | Approval | CHA, FCT, FSA | 3
FMSV2020090009 | 2G361K530 | EG-530WR | 2020-09-03 17:05:15 |  | Major | 1. Air leakege at RBS. 2. Stain & black spots on the image. 3. FST restricted 4. BSA & FSA deformed. | Decline | RBS, CHA, FCT, BSA, FSA | 4
FMSV2020100011 | 2G366K004 | EG-530NW | 2020-10-05 16:48:47 | 2020-10-08 15:56:30 | Minor | 1. Clogged AWJ. 2. U/D loosen. | Approval | AWT, OSA | 1
FMSV2021030040 | 2G366K004 | EG-530NW | 2021-03-10 14:06:34 | 2021-03-12 09:12:59 | Minor | 1. AW restricted (priority) 2. LGB broken 01 side 3. bsa defomred .Customer just want AW become normal | Approval | AWT | 2
FMSV2019090019 | 5G391K103 | EG-600WR | 2019-09-10 11:50:11 | 2019-09-11 09:51:54 | Minor |  | Approval | CHA, AWT | 1
FMSV2020100012 | 5G391K103 | EG-600WR | 2020-10-05 16:55:21 | 2020-10-07 15:33:36 | Minor | 1. Clogged AWT and WJ. 2. U/D/L/R loosen. | Approval | AWT, OSA | 2
FMSV2021030041 | 5G391K103 | EG-600WR | 2021-03-10 14:07:53 | 2021-03-12 08:11:31 | Minor | Clogged AWJ, AWT | Approval | AWT | 3
FMSV2020100059 | 6G361K178 | EG-530WR | 2020-10-26 15:02:50 | 2020-10-28 13:53:59 | Minor |  | Approval | AWT | 1
FMSV2019100043 | 6G361K179 | EG-530WR | 2019-10-28 09:33:46 | 2019-11-01 10:07:39 | Minor |  | Approval | AWT, ANGL | 1
FMSV2020100043 | 6G361K179 | EG-530WR | 2020-10-15 15:25:54 | 2020-10-22 16:53:54 | Minor | Clogged AWJ | Approval | AWT | 2
FMSV2022120005 | 6G361K179 | EG-530WR | 2022-12-05 15:59:49 |  | Major | 1. Noise on image 2. Rainbow on image 3. RC switch broken 4. BSA and FSA deformed | Decline | CHA, SW, BSA, FSA | 3