## Summary
No. | Model | Serial | Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
1 | EC-590ZW3/M | 1C610D007 | Receive | 2023-12-12 18:59
Repair
|  |  | Size | Major
|  |  | Issue | FSB, FCT, OSA
|  |  | RMA# | FMSV2023120017
|  |  | Defect | FSB damaged_x000D_ FCT leaked_x000D_ BSA loosen
|  |  | Parts | LG FLEXIBLE SECTION ASSY ROHS FCT-C283A PIPE ROHS RBS-C43A NUT ASS'Y
|  |  | Approval | Approval
2 | EC-530LP | 1C657K030 | Receive | 2019-08-01 10:57
|  |  | Repair | 2019-08-07 08:25
|  |  | Size | Minor
|  |  | Issue | VCA
|  |  | RMA# | FMSV2019080003
|  |  | Defect | VCA CALBLE BROKEN
Parts
|  |  | Approval | Approval
3 | EN-580T | 1C675K021 | Receive | 2020-11-03 16:17 | 2022-04-25 18:53 | 2022-11-14 17:51 | 2023-04-03 17:13
|  |  | Repair | 2020-11-10 10:19 | 2022-08-04 18:03 | 2022-11-16 15:34
|  |  | Size | Minor | Major | Minor | Major
|  |  | Issue | CHA | DRY, FCT | O | FCT, BSA
|  |  | RMA# | FMSV2020110008 | FMSV2022040072 | FMSV2022110035 | FMSV2023040004
|  |  | Defect | Abnormal image_x000D_ U/D/L/R were loosen | 1.VCA connector loosen._x000D_ 2.Black spot's at 1 o'clock._x000D_ 3.FCT scratch. | Blue image when scope is connected to VP7000 | 1.Leaked at FCT_x000D_ 2.BSA Stiff_x000D_ 3.Damaged glue at DTE
|  |  | Parts |  | PIPE ROHS FCT-FC675A(REPAIR) ROHS RUBBER FOR BSA ROHS CHA-FC675A-6 ROHS OSA |  | PIPE ROHS FCT-FC675A(REPAIR) ROHS B TUBE FC675A ROHS AIR/WATER TUBE ASS'Y BENDING SECTION C660A ROHS RUBBER FOR BSA ROHS SLEEVE(ANGLE)5 FC675A ROHS NUT ASS'Y
|  |  | Approval | Approval | Approval | Approval | Cancel
4 | EN-580T | 1C675K022 | Receive | 2020-10-08 17:50 | 2020-12-30 10:48
|  |  | Repair | 2020-11-03 11:08
|  |  | Size | Major | Major
|  |  | Issue | BSA | DRY, BSA, FCT
|  |  | RMA# | FMSV2020100024 | FMSV2020120080
|  |  | Defect | BSA up was broken | 1. Bending Left broken._x000D_ 2. Image black spot at 1 o'clock._x000D_ 3. FCT leak and block.
|  |  | Parts | BENDING SECTION C660A ROHS RUBBER FOR BSA ROHS SLEEVE(ANGLE)5 FC675A ROHS NUT ASSY CHA-FC675A-6 ROHS 580 UNIFIED AFE PCB MICON PCB FG388A-6 ROHS | BENDING SECTION C660A ROHS RUBBER FOR BSA ROHS SLEEVE(ANGLE)5 FC675A ROHS NUT ASSY FCT-FC675A(REPAIR) ROHS PIPE ROHS AIR/WATER TUBE ASS'Y SLEEVE FC660A ROHS SCREW B TUBE FC675A ROHS
|  |  | Approval | Approval | Cancel
5 | EC-600WI | 1C692K117 | Receive | 2019-02-26 16:18
Repair
|  |  | Size | Major
|  |  | Issue | CHA, FCT
|  |  | RMA# | FMSV2018050055
|  |  | Defect | Blue image, black spot at 7 o'clock_x000D_ FCT was restricted
|  |  | Parts | RUBBER FOR BSA ROHS CAMERA HEAD ASS'Y RING FC645A-5 ROHS FCT ASSY FC692A ROHS
|  |  | Approval | Cancel
6 | EC-600WI | 1C692K118 | Receive | 2018-12-07 10:56 | 2019-02-28 09:50
Repair
|  |  | Size | Other | Major
|  |  | Issue | O | CHA, LGB, BSA, AWT
|  |  | RMA# | FMSV2018050057 | FMSV2019020016
|  |  | Defect |  | Black spot on image_x000D_ Air water jet clogged_x000D_ Wires lossen_x000D_ LGB broken about 20%
|  |  | Parts |  | LG FC692A ROHS RUBBER FOR BSA ROHS CAMERA HEAD ASS'Y
|  |  | Approval | NoFault | Cancel
7 | EC-600WL | 1C692K141 | Receive | 2019-07-26 09:52 | 2021-03-11 11:02 | 2021-06-29 11:46 | 2021-11-08 15:58 | 2023-05-29 18:49 | 2023-07-31 12:31
|  |  | Repair | 2020-03-31 16:33 | 2021-03-16 15:28 | 2021-07-01 10:02 | 2021-12-21 17:53
|  |  | Size | Major | Minor | Minor | Major | Minor | Major
|  |  | Issue | CHA, SW | DWA | SW | FSA | CLN | FSB, VCA
|  |  | RMA# | FMSV2019070047 | FMSV2021030049 | FMSV2021060048 | FMSV2021110032 | FMSV2023050073 | FMSV2023070076
|  |  | Defect | Leaked at FR switch_x000D_ CHA broken cable | DW BROKEN | leaked at MM switch |  |  | 1.Leaked at FSB_x000D_ 2.VCA connector worn out
|  |  | Parts | FR SWITCH RUBBER FOR BSA ROHS CAMERA HEAD ASSY FC688A ROHS |  | MM SWITCH | ISA RING 2 ISA RING 1 FLEXIBLE SECTION ASSY RUBBER FOR BSA ROHS BELT |  | LG FLEXIBLE SECTION ASSY ROHS VIDEO CONNECTOR ASS'Y BELT
|  |  | Approval | Approval | Approval | Approval | Approval | Inspection | Cancel
8 | EC-600WI | 1C692K142 | Receive | 2016-08-05 17:08 | 2018-05-10 13:38 | 2019-10-28 16:09
|  |  | Repair | 2016-10-12 13:58
|  |  | Size | Major | Major | Major
|  |  | Issue | ISA | BSA | BSA, AWT
|  |  | RMA# | FMSV2016080010 | FMSV2018050006 | FMSV2019100055
|  |  | Defect | Black spots at 7 o'clock | UP wire is broken at BSA | BSA deformed_x000D_ AWT clogged
|  |  | Parts | INSERTING SECTION NUT ASS''Y SLEEVE | INSERTING SECTION NUT ASSY ROHS SLEEVE ROHS WIRE ASSY ROHS | BENDING SECTION ASSY RUBBER FOR BSA ROHS SLEEVE ROHS NUT ASSY ROHS RING FC645A-5 ROHS FCT ASSY FC692A ROHS AW-TUBE FC644A-5 ROHS WJ TUBE FC644A-5 ROHS
|  |  | Approval | Approval | Transfer | Cancel
9 | EC-600WL | 1C692K159 | Receive | 2019-07-23 11:06 | 2020-11-03 16:28
|  |  | Repair | 2019-10-08 13:51
|  |  | Size | Minor | Major
|  |  | Issue | LGB, DWA | CHA, FSA, VCA
|  |  | RMA# | FMSV2019070042 | FMSV2020110016
|  |  | Defect | U/D BROKEN, LG LOOSEN | 1. Image failed at 5 o'clock._x000D_ 2. FSA worn out._x000D_ 3. VCA connector leaked and loosen.
|  |  | Parts | WIRE ASSY ROHS WIRE ASSY ROHS | CHA-C688A RUBBER FOR BSA ROHS FLEXIBLE SECTION ASSY SLIDE SLEEVE ISA RING 2 FSA RING O-RING
|  |  | Approval | Approval | Cancel
10 | EC-600WI | 1C692K651 | Receive | 2018-05-24 10:42 | 2019-05-21 10:42 | 2019-08-12 15:17 | 2020-08-18 13:16 | 2021-03-08 10:35 | 2021-04-14 11:33
|  |  | Repair | 2018-10-10 09:38 | 2019-06-20 16:30 | 2019-08-27 09:34 | 2020-08-21 10:16 | 2021-03-08 17:41
|  |  | Size | Major | Minor | Minor | Minor | Minor | Maintain
|  |  | Issue | BSA, FSA, RBS | DWA, FSA, VCA | VCA | BSA | VCA | E
|  |  | RMA# | FMSV2018050051 | FMSV2019050033 | FMSV2019080055 | FMSV2020080040 | FMSV2021030026 | FMSV2021040040
|  |  | Defect |  | FSA Leakage, UP/DOWN broken, slide sleeve was damaged |  | UP/ LEFT WAS LOOSEN |  | not received
|  |  | Parts | BENDING SECTION ASSY RUBBER FOR BSA ROHS RING FC645A-5 ROHS FCT ASSY FC692A ROHS AW-TUBE FC644A-5 ROHS WJ TUBE FC644A-5 ROHS SLEEVE ROHS NUT ASSY ROHS |  |  | NUT ASSY
|  |  | Approval | Approval | Approval | Approval | Approval | Approval | Decline
11 | EC-600WI | 1C692K802 | Receive | 2019-11-11 07:47 | 2020-10-06 15:32 | 2020-12-30 10:47 | 2021-04-08 14:23 | 2021-11-03 16:55 | 2022-04-14 18:00
|  |  | Repair |  | 2020-10-07 16:58 |  | 2021-04-29 16:54 | 2021-11-04 16:17
|  |  | Size | Major | Minor | Major | Major | Minor | Major
|  |  | Issue | BSA, FSA | DRY, O | FSA | FSA, BSA | OSA | FSB
|  |  | RMA# | FMSV2019110025 | FMSV2020100014 | FMSV2020120079 | FMSV2021040024 | FMSV2021110012 | FMSV2022040027
|  |  | Defect | 1. FSA leaking_x000D_ 2. BSA deformed | 1. Foggy on image_x000D_ 2. FSA leaking _x000D_ 3. BSA deformed |  | FSA scratched_x000D_ BSA damaged | U/D loose | FSB broken_x000D_ BSA wire loosen._x000D_
|  |  | Parts | BENDING SECTION ASSY RUBBER FOR BSA ROHS SCREW SCREW FSA-C692A ROHS ISA RING 2 O-RING FSA RING SLEEVE ROHS NUT ASSY ROHS | RUBBER FOR BSA ROHS | ISA RING 2 ISA RING 1 FLEXIBLE SECTION ASSY RUBBER FOR BSA ROHS | ISA RING 2 ISA RING 1 FLEXIBLE SECTION ASSY BENDING SECTION ASSY RUBBER FOR BSA ROHS AW TUBE ASS'Y WATER JET TUBE PIPE SLEEVE NUT ASSY |  | LG FLEXIBLE SECTION ASSY ROHS
|  |  | Approval | Cancel | Approval | Cancel | Approval | Approval | Transfer
12 | EC-600WI | 1C692K803 | Receive | 2020-03-26 15:50 | 2020-08-31 15:55 | 2020-11-09 11:10 | 2020-11-19 16:29 | 2021-01-07 15:58 | 2021-10-07 13:48 | 2021-10-25 18:06
|  |  | Repair | 2020-03-26 16:16 | 2020-09-08 17:42 | 2020-11-10 08:56 | 2020-12-08 17:15 | 2021-01-13 14:49 | 2021-10-22 15:27 | 2021-12-09 17:06
|  |  | Size | Minor | Minor | Minor | Minor | Minor | Major | Major
|  |  | Issue | OSA | NOZ, OSA | OSA | RPR, OSA | OSA | FSA | CHA
|  |  | RMA# | FMSV2020030054 | FMSV2020080065 | FMSV2020110040 | FMSV2020110064 | FMSV2021010011 | FMSV2021100013 | FMSV2021100043
|  |  | Defect | U/D loosen | Nozzle blocked_x000D_ U/D loosen | angle control UP Down loosen |  | R-L angle loosen | FSA leak | Blue Image
|  |  | Parts |  | Nozzle |  |  |  | ISA RING 2 ISA RING 1 FLEXIBLE SECTION ASSY RUBBER FOR BSA ROHS | CHA-C688A RUBBER FOR BSA ROHS
|  |  | Approval | Approval | Approval | Approval | Approval | Approval | Approval | Approval
13 | EC-600WI | 1C692K806 | Receive | 2019-06-14 11:11 | 2019-07-15 16:10 | 2020-07-23 10:58 | 2020-11-03 17:01
|  |  | Repair | 2019-06-18 13:54 | 2020-01-10 15:57 | 2020-07-28 08:12
|  |  | Size | Minor | Major | Major | Major
|  |  | Issue | CLN | LGB | ANGL | FSA, BSA
|  |  | RMA# | FMSV2019060020 | FMSV2019070026 | FMSV2020070059 | FMSV2020110033
|  |  | Defect | Leaked at O ring U/D | Broken light guide 50% |  | 1. FCT closed_x000D_ 2. BSA deformed
|  |  | Parts |  | CHA-C688A Nozzle CAP(Old DTA) RING FC645A-5 ROHS FCT ASSY FC692A ROHS AW-TUBE FC644A-5 ROHS WJ TUBE FC644A-5 ROHS LG FC692A ROHS BENDING SECTION ASSY FSA-C692A ROHS ISA RING 2 O-RING FSA RING RUBBER FOR BSA ROHS SLEEVE ROHS NUT ASSY ROHS |  | BENDING SECTION ASSY RUBBER FOR BSA ROHS PIPE SLEEVE NUT ASSY FCT PIPE FCT ASSY FC692A ROHS AW TUBE ASS'Y WATER JET TUBE NOZZLE CAP
|  |  | Approval | Approval | Approval | Approval | Cancel
14 | EC-600WI | 1C692K826 | Receive | 2020-01-15 14:11 | 2020-10-06 15:29 | 2021-03-17 10:59 | 2021-10-12 16:31 | 2022-04-13 18:58
|  |  | Repair |  | 2020-10-07 16:59 | 2021-03-31 15:59 | 2021-12-13 10:44
|  |  | Size | Major | Minor | Major | Major | Maintain
|  |  | Issue | CHA, DWA, RBS | DRY, RBS, DWA | BSA, FSA, WGA/WGB | FCT, DRY, VCA | O
|  |  | RMA# | FMSV2020010023 | FMSV2020100013 | FMSV2021030069 | FMSV2021100021 | FMSV2022040023
|  |  | Defect | 1) Black spots on the image at 2 o'clock. _x000D_ 2) Air leaked at RBS. _x000D_ 3) Drum wire UP is broken. _x000D_ 4) Net of BSA broken | 1) Black spots on the image at 2 o'clock. _x000D_ 2) Air leaked at RBS. _x000D_ 3) Drum wire UP is broken. _x000D_ 4) Net of BSA broken | 1. Bending UP broken._x000D_ 2. FSA worn out._x000D_ 3. Lack glue at LG lens. |  | Inspection OK
|  |  | Parts | CHA-C688A AW-TUBE FC644A-5 ROHS WJ TUBE FC644A-5 ROHS Nozzle CAP(New DTA) BENDING SECTION ASSY SLEEVE ROHS NUT ASSY ROHS RUBBER FOR BSA ROHS WIRE ASSY ROHS WIRE ASSY ROHS | WIRE ASSY ROHS WIRE ASSY ROHS RUBBER FOR BSA ROHS | FSA-C692A ROHS BENDING SECTION ASSY RUBBER FOR BSA ROHS ISA RING 2 O-RING ISA RING 1 PIPE SLEEVE NUT ASSY | FCT RING FCT-C692A AW TUBE ASS'Y WATER JET TUBE RUBBER FOR BSA ROHS
|  |  | Approval | Decline | Approval | Approval | Approval | Transfer
15 | EC-600WI | 1C692K833 | Receive | 2020-07-01 14:39 | 2020-07-23 16:49 | 2020-10-15 10:36 | 2021-02-25 16:46 | 2021-09-15 18:59 | 2022-02-21 13:47 | 2022-04-15 10:43
|  |  | Repair | 2020-07-02 11:51 | 2020-07-28 08:12 | 2020-10-16 13:13 | 2021-03-31 09:37 | 2021-09-17 11:52 | 2022-03-02 18:00 | 2022-04-19 15:40
|  |  | Size | Minor | Minor | Minor | Major | Minor | Major | Minor
|  |  | Issue | OSA | OSA | VCA | VCA, LGB | DWA | DRY, DEC | VCA
|  |  | RMA# | FMSV2020070008 | FMSV2020070061 | FMSV2020100037 | FMSV2021020029 | FMSV2021090029 | FMSV2022020022 | FMSV2022040030
|  |  | Defect |  | Wrong position R/L |  | 1. Image noise due to VCA cable broken_x000D_ 2. LGB has broken one side | Drum wire Up/Down broken. | black spots | Image noise due to VCA cable broken
|  |  | Parts |  |  |  | LG FC692A ROHS RUBBER FOR BSA ROHS | WIRE ASSY WIRE ASSY
|  |  | Approval | Approval | Approval | Approval | Approval | Approval | Approval | Approval
16 | EC-600WI | 1C692K835 | Receive | 2018-05-31 13:54 | 2018-09-10 16:34 | 2019-05-22 15:55 | 2020-03-17 11:05 | 2020-12-17 16:25 | 2021-03-05 16:39 | 2021-06-10 17:57
|  |  | Repair | 2018-08-28 16:03 | 2018-09-11 16:57 | 2019-08-30 10:46 | 2020-07-29 09:40 | 2020-12-18 09:18 | 2021-03-10 14:59
|  |  | Size |  | Minor | Major | Major | Major | Minor | Major
|  |  | Issue | FCT | WJT | CHA, AWT | DWA, CHA | AWT | WGB | FSA, FCT, BSA
|  |  | RMA# | FMSV2018050063 | FMSV2018090019 | FMSV2019050052 | FMSV2020030036 | FMSV2020120041 | FMSV2021030021 | FMSV2021060011
|  |  | Defect | FCT scratched. Wires are loosen |  | Leaked at Water Jet tube, Black spots at 1h clock | 1. DW broken and loosen_x000D_ 2. Black spotted at 2-o'clok | Air and water were restricted |  | 1. FSA leaked_x000D_ 2. FCT restricted_x000D_ 3. BSA stiff
|  |  | Parts | RING FC645A-5 ROHS FCT ASSY FC692A ROHS AW-TUBE FC644A-5 ROHS WJ TUBE FC644A-5 ROHS BENDING SECTION ASSY RUBBER FOR BSA ROHS SLEEVE ROHS NUT ASSY ROHS |  | AW-TUBE FC644A-5 ROHS WJ TUBE FC644A-5 ROHS RUBBER FOR BSA ROHS CAMERA HEAD ASS'Y | CHA-C688A WIRE ASSY ROHS WIRE ASSY ROHS RUBBER FOR BSA ROHS |  |  | FCT-C692A FCT RING WATER JET TUBE AW TUBE ASS'Y FLEXIBLE SECTION ASSY ISA RING 2 O-RING ISA RING 1 BENDING SECTION ASSY NUT ASSY PIPE SLEEVE RUBBER FOR BSA ROHS
|  |  | Approval | Approval | Approval | Approval | Approval | Approval | Approval | Decline
17 | EC-600WI | 1C692K836 | Receive | 2020-03-26 15:06 | 2020-09-21 10:32
|  |  | Repair | 2020-03-27 18:40 | 2020-09-23 08:12
|  |  | Size | Minor | Minor
|  |  | Issue | ANGL | ANGL
|  |  | RMA# | FMSV2020030053 | FMSV2020090048
Defect
Parts
|  |  | Approval | Approval | Approval
18 | EC-760ZP-V/M | 1C730K096 | Receive | 2019-09-16 17:09
Repair
|  |  | Size | Major
|  |  | Issue | ANGL, CHA
|  |  | RMA# | FMSV2019090030
Defect
|  |  | Parts | L/R KNOB ASSY(20N.CM) ROHS FCTINLET COVER RUBBER CHA-C730A RBS-C716A COVER L LD PCB ASSY HEAT CONDUCTION SHEET
|  |  | Approval | Decline
19 | EG-590ZW | 1G349K025 | Receive | 2019-10-08 11:09
Repair
|  |  | Size | Major
|  |  | Issue | CHA, FCT, VCA, LGB
|  |  | RMA# | FMSV2019100008
|  |  | Defect | blur and stain image, FCT restricted, VCA scratch, LGB broken 40%
|  |  | Parts | CAMERA HEAD ASSY / DRIVE SHAFT FORCEPS TUBE PIPE ROHS AIR TUBE W-TUBE FG246A ROHS RBS-C112B VIDEO CONNECTOR ASSY ROHS CAP+UV FG258A ROHS NOZZLE LGB-G258A
|  |  | Approval | Cancel
20 | EG-600WR | 1G391K219 | Receive | 2018-05-24 11:20 | 2018-12-11 14:23 | 2019-12-05 12:05
|  |  | Repair | 2018-08-27 08:47 | 2019-09-12 11:23
|  |  | Size | Major | Major | Major
|  |  | Issue | FCT, WJT | BSA | CHA, BSA
|  |  | RMA# | FMSV2018050053 | FMSV2018120017 | FMSV2019120009
|  |  | Defect | Water jet is leaked at distal end | Bending deformed
|  |  | Parts | RUBBER FOR BSA ROHS FCT ASSY FG391A ROHS FCT RING FG381A-6 ROHS TUBE(WJ) FG391A ROHS AW-TUBE FG391A ROHS | BENDING SECTION ASSY RUBBER FOR BSA ROHS RETAINER RING TAPER RING SLEEVE ROHS NUT ASSY ROHS FCT RING FG381A-6 ROHS FCT ASSY FG391A ROHS AW-TUBE FG391A ROHS TUBE(WJ) FG391A ROHS | FITTING PART FCT ASSY FG391A ROHS AIR/WATER TUBE ASS''Y TUBE(WJ) FG391A ROHS CAMERA HEAD ASS'Y BENDING SECTION ASSY RBS-G181A SLEEVE ROHS NUT ASSY ROHS
|  |  | Approval | Approval | Approval | Cancel
21 | EG-600WR | 1G391K223 | Receive | 2019-08-09 14:42
Repair
Size
|  |  | Issue | CHA, FCT, VCA, BSA
|  |  | RMA# | FMSV2019080045
|  |  | Defect | CHA len broken, Fct restriction, bsa deformed._x000D_ VCA leaking
|  |  | Parts | CHA-FG391A ROHS AW-TUBE FG391A ROHS TUBE(WJ) FG391A ROHS FCT ASSY FG391A ROHS FCT RING FG381A-6 ROHS BENDING SECTION ASSY RBS-G181A SCREW SLEEVE ROHS NUT ASSY ROHS SLEEVE ROHS VIDEO CONNECTOR ASS/Y ROHS
|  |  | Approval | Decline
22 | EG-600WR | 1G391K324 | Receive | 2019-08-09 14:40
Repair
|  |  | Size | Major
|  |  | Issue | CHA, BSA, FCT, VCA
|  |  | RMA# | FMSV2019080044
|  |  | Defect | 1. Stain image_x000D_ 2. FCT restriction_x000D_ 3. BSA deformed_x000D_ 4. VCA loosen
|  |  | Parts | CHA-FG391A ROHS AW-TUBE FG391A ROHS TUBE(WJ) FG391A ROHS FCT RING FG381A-6 ROHS FCT ASSY FG391A ROHS CAP NOZZLE BENDING SECTION ASSY SLEEVE ROHS NUT ASSY ROHS RBS-G181A
|  |  | Approval | Decline
23 | EG-600WR | 1G391K366 | Receive | 2015-09-02 14:05 | 2016-11-18 08:46 | 2016-11-28 11:48 | 2018-05-10 13:39 | 2019-01-09 11:14 | 2019-05-06 15:10
|  |  | Repair | 2015-12-11 10:52 | 2016-11-21 07:59 | 2016-12-13 15:29 | 2018-05-22 10:40 | 2019-01-25 13:50
|  |  | Size | Major | Minor | Major | Major | Major | Major
|  |  | Issue | CHA | VCA | FCT, BSA | ISA | OSA | VCA, BSA
|  |  | RMA# | FMSV2015090001L | FMSV2016110010 | FMSV2016110014 | FMSV2018050003 | FMSV2019010003 | FMSV2019050001
|  |  | Defect | Abnormal image | VCA was damaged |  | WGA Lens cracked_x000D_ AWT leaked_x000D_ | Suction valve assembly is always suction without press buuton | Up wire broken_x000D_ VCA damaged
|  |  | Parts | CAMERA HEAD ASS'Y BENDING RUBBER | VCA | BENDING RUBBER BENDING SECTION ASSY NUT ASS''Y PIPE SLEEVE FORCEPS TUBE RING | INSERTING SECTION SLEEVE ROHS NUT ASSY ROHS | SUCTION VALVE ASS'Y 500_2 | BENDING SECTION ASSY RUBBER FOR BSA ROHS SLEEVE ROHS NUT ASSY ROHS VIDEO CONNECTOR ASS/Y ROHS
|  |  | Approval | Approval | Approval | Approval | Approval | Approval | Cancel
24 | EG-600WR | 1G391K368 | Receive | 2019-08-09 14:43
Repair
|  |  | Size | Major
|  |  | Issue | CHA, FCT, FSA, SW, BSA
|  |  | RMA# | FMSV2019080046
|  |  | Defect | 1. White spots and stain image at 7 o'clock._x000D_ 2. FCT is restriction at bending section._x000D_ 3. FSA is deformed._x000D_ 4. Leaking at MM Switch._x000D_ 5. BSA is deformed._x000D_
|  |  | Parts | INSERTING SECTION SLEEVE ROHS NUT ASSY ROHS MM SWITCH
|  |  | Approval | Decline
25 | EG-600WR | 1G391K420 | Receive | 2015-11-05 11:33 | 2016-06-10 15:58 | 2016-08-05 16:46 | 2018-05-10 14:16 | 2018-11-26 15:30
|  |  | Repair | 2016-06-08 18:43 | 2016-06-10 16:02 | 2016-10-12 13:33
|  |  | Size | Major | Other | Major | Minor | Major
|  |  | Issue | ISA, VCA |  | ISA | DWA | LPA, BSA, VCA
|  |  | RMA# | FMSV2015110006L | FMSV2016060010 | FMSV2016080008 | FMSV2018050007 | FMSV2018110059
|  |  | Defect | 1) Black dots appeared on the endoscopic image at 7 o'clock position_x000D_ 2) VCA was losen because the slide sleeve was deformed | OK | Black spot at 7o'clock
|  |  | Parts | PIPE SLEEVE NUT ASS''Y INSERTING SECTION SLIDE SLEEVE |  | INSERTING SECTION NUT ASS''Y PIPE SLEEVE |  | BELT AIRTIGHT CAP ROHS BENDING SECTION ASSY RUBBER FOR BSA ROHS SLEEVE ROHS NUT ASSY ROHS WIRE ASSY(U) DWA-670 ROHS WIRE ASSY(D) DWA-615 ROHS
|  |  | Approval | Approval | Approval | Approval | Cancel | Cancel
26 | EG-760R | 1G402K918 | Receive | 2019-07-19 15:13 | 2020-09-07 15:09 | 2021-09-07 17:58 | 2021-11-15 14:00 | 2022-01-25 17:08
|  |  | Repair | 2019-08-16 15:08 |  | 2021-09-17 11:49 | 2021-12-09 17:05 | 2022-02-17 15:06
|  |  | Size | Minor | Major | Minor | Major | Major
|  |  | Issue | CHA | FSB, FSA | OSA | FSB, OSA | FCT, DRY, FSA
|  |  | RMA# | FMSV2019070033 | FMSV2020090014 | FMSV2021090011 | FMSV2021110055 | FMSV2022010037
|  |  | Defect |  | FSB broken_x000D_ FSA deformed | U/D loosen |  | leaked test, CHA black spot
|  |  | Parts | SC600 FPGA PCB ASSY SC600 MCU PCB ASSY | FSA-G402A RBS-G181A FSB A FCTINLET COVER RUBBER |  | FSB A | NOZZLE J TUBE A/W TUBE BSA-G400A ROUND CORNER SLEEVE FCTINLET COVER RUBBER HANGING SLEEVE WIRE SLEEVE CAP RBS-G181A
|  |  | Approval | Approval | Cancel | Approval | Approval | Approval
27 | EG-760Z | 1G403K346 | Receive | 2020-05-13 10:29 | 2020-10-01 13:28 | 2021-01-18 09:30
|  |  | Repair | 2020-05-18 10:30 | 2020-10-01 17:35
|  |  | Size | Major | Minor | Major
|  |  | Issue | WGB, SW, FCT | ANGL | AWT, ISA, LC
|  |  | RMA# | FMSV2020050028 | FMSV2020100003 | FMSV2021010033
|  |  | Defect |  |  | noise on image_x000D_ scope leak_x000D_
|  |  | Parts | SW BOX ASSY (5 BUTTON) FCT-G402A FITTING PART CAP LG LENS ASSY FG378A-6 ROHS NOZZLE FCTINLET COVER RUBBER HANGING SLEEVE WIRE SLEEVE FSA-G403A ROUND CORNER SLEEVE BENDING RUBBER BSA-G403A | SLIDEING RING O RING W2 | ISA-G403A HANGING SLEEVE WIRE SLEEVE FCTINLET COVER RUBBER LG COVER MIDDLE ASSY
|  |  | Approval | Approval | Approval | Decline
28 | XL-4450 | 1S094K191 | Receive | 2022-02-28 10:18 | 2023-08-21 17:26
|  |  | Repair | 2022-03-08 12:09
|  |  | Size | Other | Other
|  |  | Issue | O | IRS
|  |  | RMA# | FMSV2022020038 | FMSV2023080081
Defect
|  |  | Parts |  | IRIS ASSY
|  |  | Approval | Approval | Cancel
29 | XL-4450 | 1S094K192 | Receive | 2019-12-16 10:17 | 2020-08-12 10:34 | 2022-09-20 17:45 | 2023-01-03 13:49 | 2023-02-02 18:13
|  |  | Repair |  | 2020-08-20 11:25 | 2022-09-26 12:08 | 2023-01-04 17:04 | 2023-04-12 10:27
|  |  | Size | Other | Other | Minor | Other | Other
|  |  | Issue | O | O, PCB | O | O | IRS
|  |  | RMA# | FMSV2019120019 | FMSV2020080026 | FMSV2022090043 | FMSV2023010005 | FMSV2023020009
|  |  | Defect | 1) The pump is dirty so air come out weak ( cleaned --> Pump is OK) _x000D_ 2) 02 Springs for LG Guide are broken. | 1) The pump is dirty so air come out weak ( cleaned --> Pump is OK). _x000D_ 2) 02 Springs for LG Guide are broken. _x000D_ 3) The light coming out of the hole is always largest ---> Can not control the Iris unit due to LIGHT CONTROL _x000D_ PCB is broken (Dirt and mold _x000D_ appear on the board). |  |  | Iris failure
|  |  | Parts | SPRING LIGHT CONTROL PCB |  |  |  | IRIS ASSY
|  |  | Approval | Cancel | Approval | Approval | Approval | Approval
30 | XL-4450 | 1S094K193 | Receive | 2021-02-22 17:25 | 2021-03-16 13:36 | 2022-04-20 13:57
|  |  | Repair | 2021-02-25 16:23 |  | 2022-04-20 19:42
|  |  | Size | Other | Other | Other
|  |  | Issue | ESS Replace Packing/Ring | IRS, ESS Replace Packing/Ring | IRS, O
|  |  | RMA# | FMSV2021020023 | FMSV2021030060 | FMSV2022040042
|  |  | Defect | error connected between scopes with XL-4450 | 1. Movements failure_x000D_ 2. Wear air | 1. Movements failure_x000D_ 2. Wear air
|  |  | Parts |  | IRIS ASSY RING PACKING ISOLATION PLATE ISOLATION RING FITTING PART FITTING PART LG GUIDE HOLDER | IRIS ASSY RING PACKING ISOLATION PLATE ISOLATION RING FITTING PART FITTING PART LG GUIDE HOLDER
|  |  | Approval | Approval | Cancel | Approval
31 | XL-4450 | 1S094K194 | Receive | 2021-10-26 17:51 | 2022-06-07 10:34 | 2022-11-08 18:46
|  |  | Repair | 2021-10-27 15:13 | 2022-06-09 12:05 | 2022-11-10 19:32
|  |  | Size | Other | Minor | Minor
|  |  | Issue | O | O | PUM
|  |  | RMA# | FMSV2021100051 | FMSV2022060021 | FMSV2022110016
|  |  | Defect |  | weak air_x000D_ | Air weak_x000D_
Parts
|  |  | Approval | Approval | Approval | Approval
32 | XL-4450 | 1S094K195 | Receive | 2021-02-23 11:01 | 2021-10-14 19:12 | 2022-11-08 18:46 | 2023-02-02 18:12
|  |  | Repair | 2021-02-24 08:32 | 2021-10-15 15:26 | 2022-11-17 19:03 | 2023-04-12 10:27
|  |  | Size | Other | Other | Minor | Other
|  |  | Issue | PUM | PUM | O | IRS
|  |  | RMA# | FMSV2021020024 | FMSV2021100026 | FMSV2022110017 | FMSV2023020007
|  |  | Defect |  |  |  | Iris failure
|  |  | Parts |  |  |  | IRIS ASSY
|  |  | Approval | Approval | Approval | Approval | Approval
33 | VP-4450HD | 1V567K309 | Receive | 2022-02-17 11:15
Repair
|  |  | Size | Other
|  |  | Issue | PCB
|  |  | RMA# | FMSV2022020014
Defect
|  |  | Parts | MAIN PCB HEAT SINK ASSY
|  |  | Approval | Cancel
34 | VP-4450HD | 1V567K312 | Receive | 2023-01-03 13:48
Repair
|  |  | Size | Other
|  |  | Issue | PCB
|  |  | RMA# | FMSV2023010004
Defect
|  |  | Parts | MAIN PCB HEAT SINK ASSY
|  |  | Approval | Decline
35 | EC-590ZW3/L | 2C667D037 | Receive | 2021-03-11 10:35
|  |  | Repair | 2021-03-18 17:11
|  |  | Size | Minor
|  |  | Issue | BSA, AWT
|  |  | RMA# | FMSV2021030046
|  |  | Defect | 1. bsa loosen_x000D_ 2. awj restricted_x000D_ 3. Glue of camera len broken
Parts
|  |  | Approval | Approval
36 | EC-760R-V/I | 2C728K101 | Receive | 2020-08-14 09:40 | 2020-10-01 09:47 | 2021-03-04 17:20 | 2021-11-04 13:01 | 2022-02-24 13:00
|  |  | Repair | 2020-08-14 15:27 | 2020-10-01 13:27 | 2021-05-06 08:33 |  | 2022-02-25 13:16
|  |  | Size | Minor | Minor | Major | Maintain | Minor
|  |  | Issue | AWT | RPR | RPR, DRY | O | LC
|  |  | RMA# | FMSV2020080031 | FMSV2020100001 | FMSV2021030016 | FMSV2021110018 | FMSV2022020030
|  |  | Defect |  | Leaked test | Heavy leaked
|  |  | Parts |  | O RING W2 SLIDEING RING | FCTINLET COVER RUBBER COVER L RUBBER FOR BSA ROHS LD PCB ASSY
|  |  | Approval | Approval | Approval | Approval | Inspection | Approval
37 | EC-760R-V/I | 2C728K102 | Receive | 2020-06-23 16:17 | 2020-09-08 16:52 | 2021-01-12 15:16 | 2021-04-15 16:02 | 2021-11-03 16:51 | 2021-12-30 12:41 | 2022-04-25 19:03 | 2023-09-25 18:35
|  |  | Repair |  | 2020-09-10 15:05 | 2021-01-19 09:41 | 2021-04-29 08:19 | 2021-11-04 14:43 | 2022-01-04 18:57 | 2022-04-28 11:41 | 2023-09-28 12:09
|  |  | Size | Maintain | Minor | Minor | Major | Minor | Minor | Minor | Minor
|  |  | Issue | O | LC | ANGL, OSA | WJT, OSA, DEC | WGA | OSA | CLN, BSA | BSA
|  |  | RMA# | FMSV2020060063 | FMSV2020090017 | FMSV2021010023 | FMSV2021040045 | FMSV2021110011 | FMSV2021120070 | FMSV2022040073 | FMSV2023090057
|  |  | Defect |  | Ventilation connector heavy toque | 1. Leaked at Knob_x000D_ 2. U/D Loosen | Damaged glue_x000D_ U/D loosen | Damaged glue | U/D loosen.
|  |  | Parts |  |  |  | J TUBE CAP NOZZLE RUBBER FOR BSA ROHS
|  |  | Approval | Inspection | Approval | Approval | Approval | Approval | Approval | Approval | Approval
38 | EC-760ZP-V/M | 2C730K032 | Receive | 2019-08-30 10:12 | 2020-08-18 11:15 | 2021-09-07 17:59 | 2022-03-18 20:09 | 2022-07-27 13:41 | 2023-02-08 18:54
|  |  | Repair | 2020-04-29 07:50 | 2020-08-19 12:46 | 2021-09-17 11:50
|  |  | Size | Major | Minor | Minor | Major | Other | Major
|  |  | Issue | CHA | CAP | OSA | BSA, LGB | INS | LGB, BSA, FSB
|  |  | RMA# | FMSV2019080104 | FMSV2020080039 | FMSV2021090012 | FMSV2022030039 | FMSV2022070098 | FMSV2023020022
|  |  | Defect | Abnormal Image, some time blue image |  | U/D lossen | 1.BSA deformed_x000D_ 2.LGB broken | LGB broken one side_x000D_ | 1.LGB broken_x000D_ 2.BSA deformed_x000D_ 3.FSB worn out_x000D_ 4.Damaged glue at DEC
|  |  | Parts | CHA-C730A RBS-C716A COVER L FCTINLET COVER RUBBER |  |  | BSA-C727A RBS ROUND CORNER SLEEVE FCT RING FCT-C730A A/W TUBE J TUBE LGB-C727A FCTINLET COVER RUBBER COVER L |  | LGB-C727A FCT RING FCT-C730A A/W TUBE J TUBE BSA-C727A RBS ROUND CORNER SLEEVE HANGING SLEEVE WIRE SLEEVE FSB A FCTINLET COVER RUBBER COVER L
|  |  | Approval | Approval | Approval | Approval | Cancel | Inspection | Cancel
39 | EG-600WR | 2G391K487 | Receive | 2018-05-24 11:18 | 2019-07-29 16:30 | 2019-10-16 11:22 | 2019-11-06 10:29 | 2022-01-27 19:06 | 2023-09-08 15:01
|  |  | Repair | 2018-12-12 15:16 | 2019-09-25 09:18 | 2019-10-18 10:50 |  | 2022-01-28 17:44
|  |  | Size |  | Major | Minor | Major | Major | Major
|  |  | Issue | FCT | LGB, FCT | ANGL | CHA, WJT, BSA, FSA | CHA, WJT, BSA, FSA | BSA
|  |  | RMA# | FMSV2018050052 | FMSV2019070053 | FMSV2019100019 | FMSV2019110012 | FMSV2022010042 | FMSV2023090015
|  |  | Defect | FCT restriction |  |  | 1) Black spots appears on the image. _x000D_ 2) Air leaking at Water jet tube. _x000D_ 3) BSA is deformed when bending Down. _x000D_ 4) FSA is demaged about 30 cm | Same Job FMSV2019110012 | BSA heavy torque
|  |  | Parts | FCT RING FG381A-6 ROHS FCT ASSY FG391A ROHS RUBBER FOR BSA ROHS | FCT RING FG381A-6 ROHS FCT ASSY FG391A ROHS AW-TUBE FG391A ROHS TUBE(WJ) FG391A ROHS LG-FG391A ROHS RUBBER FOR BSA ROHS |  | CAMERA HEAD ASS'Y FCT ASSY FG391A ROHS FITTING PART AW-TUBE FG391A ROHS TUBE(WJ) FG391A ROHS FSA-G391A ROHS RING FG378A-6 ROHS FSA RING BENDING SECTION ASSY SLEEVE ROHS NUT ASSY ROHS RBS-G181A | FCT ASSY FG391A ROHS FITTING PART AW-TUBE FG391A ROHS TUBE(WJ) FG391A ROHS FSA-G391A ROHS RING FG378A-6 ROHS FSA RING BENDING SECTION ASSY SLEEVE ROHS NUT ASSY RBS-G181A | BSA G391A RBS-G181A NUT ASS'Y SLEEVE ROHS FCT RING FCT-G391A AIR/WATER TUBE ASS''Y J TUBE
|  |  | Approval | Approval | Approval | Approval | Cancel | Approval | Approval
40 | EG-600WR | 2G391K630 | Receive | 2018-01-31 09:19 | 2020-05-08 13:12 | 2020-06-09 14:01 | 2021-04-05 13:37 | 2022-01-18 11:20 | 2022-11-02 11:30 | 2023-05-18 15:10 | 2023-09-27 15:48 | 2023-10-16 18:25 | 2023-12-28 10:30
|  |  | Repair |  |  |  | 2021-06-08 16:19 | 2022-01-20 11:18 |  |  | 2023-10-02 18:29
|  |  | Size | Major | Maintain | Maintain | Major | Minor | Maintain | Maintain | Minor | Maintain | Major
|  |  | Issue | ISA, CHA, WGB, BSA, FSA, FCT | O | INS | CHA, BSA, SW, DEC | SW, OSA | O | O | VCA | O | FSA, DRY, PCB
|  |  | RMA# | FMSV2018010019 | FMSV2020050015 | FMSV2020060029 | FMSV2021040011 | FMSV2022010031 | FMSV2022100083 | FMSV2023050008 | FMSV2023090061 | FMSV2023100025 | FMSV2023120050
|  |  | Defect | 1. Image has black spots, rainbow_x000D_ 2. OBL, WGB has cracked_x000D_ 3. BSA was deformed._x000D_ 4. FSA has scratched_x000D_ 5. FCT has restriction |  |  | black spots image_x000D_ BSA deformed_x000D_ | Scope SW no function_x000D_ U/D wire loosen | inspection OK |  | Noise image |  | 1. FSB broken => leaked_x000D_ 2. Noise Image => drying => CHA OK_x000D_ 3. PCB damaged_x000D_ _x000D_
|  |  | Parts | SLEEVE ROHS NUT ASSY ROHS ISA-G391A ROHS |  |  | NUT ASSY SLEEVE ROHS BSA G391A RBS-G181A NOZZLE CAP FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE SLIDE SLEEVE | [5 empty] | RBS-G181A S600 CTL PCB LG FLEXIBLE SECTION ASSY ROHS
|  |  | Approval | Transfer | Inspection | Inspection | Approval | Approval | Inspection | Inspection | Approval | Inspection | Transfer
41 | EG-600WR | 2G391K955 | Receive | 2019-08-05 13:07 | 2020-01-20 16:04 | 2020-07-30 14:55 | 2020-08-24 15:05 | 2021-10-22 15:12
|  |  | Repair | 2019-08-29 15:28 | 2020-04-10 08:21 |  | 2020-11-27 14:43
|  |  | Size | Minor | Major | Maintain | Major | Major
|  |  | Issue | RBS | CHA, OSA | O | FCT, DEC | FSA, WJT, DRY, CHA
|  |  | RMA# | FMSV2019080011 | FMSV2020010032 | FMSV2020070074 | FMSV2020080047 | FMSV2021100039
|  |  | Defect | RBS broken. | Black spots at 7 o'clock_x000D_ BSA wires loosen | Scope OK | FCT RETRICTED | FSA crrosion_x000D_ WJT leaked_x000D_ CHA has black spots
|  |  | Parts | RBS-G181A | RBS-G181A CAMERA HEAD ASS'Y |  | RBS-G181A FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE | ISA RING 2 ISA RING 1 FSA-G391A ROHS BSA G391A RING2 FG391A ROHS AIR/WATER TUBE ASS''Y J TUBE SLEEVE ROHS NUT ASSY
|  |  | Approval | Approval | Approval | Inspection | Approval | Cancel
42 | EG-600WR | 2G391K958 | Receive | 2020-02-19 13:28 | 2020-07-21 17:00 | 2020-09-29 09:08 | 2020-12-22 13:48 | 2021-04-27 15:16 | 2021-05-07 16:00
|  |  | Repair | 2020-07-16 17:02 | 2020-07-23 17:10 | 2020-09-30 14:24 | 2021-01-29 08:43 | 2021-04-29 09:24 | 2021-05-13 18:03
|  |  | Size | Major | Minor | Minor | Major | Minor | Minor
|  |  | Issue | SW, BSA, ISA | OSA | VCA | CHA, BSA, FCT | ANGL | BSA
|  |  | RMA# | FMSV2020020025 | FMSV2020070054 | FMSV2020090062 | FMSV2020120048 | FMSV2021040085 | FMSV2021050014
|  |  | Defect | RC switch is broken_x000D_ BSA wires are loosen_x000D_ Black spots at 7 o'clock | Loaner in 3 days, used case 179 cases |  | 1. Image noise and black dots_x000D_ 2. FCT was restricted_x000D_ 3. BSA was deformed
|  |  | Parts | INSERTING SECTION SLEEVE ROHS NUT ASSY ROHS RC SWITCH S600 CTL PCB SLIDE SLEEVE |  |  | FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE CAMERA HEAD ASS'Y BSA G391A RBS-G181A RING2 FG391A ROHS SLEEVE ROHS NUT ASSY
|  |  | Approval | Approval | Approval | Approval | Approval | Approval | Approval
43 | EG-600WR | 2G391K960 | Receive | 2020-03-04 11:02 | 2020-11-02 08:35 | 2021-02-26 13:28 | 2021-11-09 16:14 | 2022-04-14 10:46
|  |  | Repair | 2020-03-04 11:20 | 2020-11-20 08:53 | 2021-05-06 15:36 |  | 2022-04-18 12:05
|  |  | Size | Minor | Major | Major | Major | Major
|  |  | Issue | LPA | FCT, CHA | LC | RBS, DRY, CHA, BSA | RBS, DRY, CHA, BSA
|  |  | RMA# | FMSV2020030003 | FMSV2020110004 | FMSV2021020031 | FMSV2021110035 | FMSV2022040025
|  |  | Defect |  | Black spot and FCT block |  | RBS leak_x000D_ CHA black spots_x000D_ BSA worn out | RBS leak_x000D_ CHA black spots_x000D_ BSA worn out
|  |  | Parts |  | CAMERA HEAD ASS'Y FITTING PART FCT-G391A RBS-G181A | LG CONNECTOR LG PROBE2 FG378A-6 ROHS ROD (1.95) ROHS SPACER ROHS | BSA G391A RBS-G181A FSA-G391A ROHS ISA RING 2 ISA RING 1 J TUBE AW-TUBE FG391A ROHS SLEEVE ROHS NUT ASSY | SLEEVE ROHS NUT ASSY INSERTING SECTION VIDEO CONNECTOR ASS'Y
|  |  | Approval | Approval | Approval | Approval | Decline | Approval
44 | EG-600WR | 2G391K963 | Receive | 2020-03-02 09:52 | 2020-12-17 15:49 | 2021-06-11 08:56 | 2021-09-20 18:58 | 2021-11-12 11:05 | 2022-03-17 18:00
|  |  | Repair | 2020-04-10 14:44 |  | 2021-06-30 13:33 | 2021-09-23 18:23 |  | 2022-03-23 13:36
|  |  | Size | Major | Major | Minor | Minor | Major | Major
|  |  | Issue | CHA, FCT | ISA, FSB | VCA | SW | BSA, CLN, FSB, VCA | BSA, CLN, FSB, VCA
|  |  | RMA# | FMSV2020030001 | FMSV2020120036 | FMSV2021060015 | FMSV2021090031 | FMSV2021110040 | FMSV2022030032
|  |  | Defect | stain image at 12 o'clcok, black spots at 7 o'clock, , u/d loosen, scratch FCT, no glue on WGB | Black spots image_x000D_ FSB deformed | VCA cable broken | Leak at RC switch. | 1. BSA broken_x000D_ 2. CHA lens scratched_x000D_ 3. FSB broken_x000D_ 4. Belt of VCA broken_x000D_ 5. Noise image due to VCA cable broken | 1. BSA broken_x000D_ 2. CHA lens scratched_x000D_ 3. FSB broken_x000D_ 4. Belt of VCA broken_x000D_ 5. Noise image due to VCA cable broken
|  |  | Parts | CAMERA HEAD ASS'Y FITTING PART FCT ASSY FG391A ROHS AIR/WATER TUBE ASS''Y TUBE(WJ) FG391A ROHS RBS-G181A | BSA G391A RBS-G181A NUT ASSY SLEEVE ROHS LG FLEXIBLE SECTION ASSY ROHS |  | RC SWITCH | BSA G391A SLEEVE ROHS NUT ASSY LG FLEXIBLE SECTION ASSY ROHS RBS-G181A BELT | BSA G391A SLEEVE ROHS NUT ASSY LG FLEXIBLE SECTION ASSY ROHS RBS-G181A BELT
|  |  | Approval | Approval | Cancel | Approval | Approval | Decline | Approval
45 | EG-600WR | 2G391K967 | Receive | 2020-11-03 16:33 | 2021-04-01 16:33 | 2021-10-15 15:30
|  |  | Repair |  | 2021-04-12 08:33 | 2022-02-11 11:06
|  |  | Size | Major | Major | Major
|  |  | Issue | CHA, FCT, FSA | CHA, FCT, FSA, VCA | SW, BSA, FCT
|  |  | RMA# | FMSV2020110022 | FMSV2021040004 | FMSV2021100030
|  |  | Defect | 1. Image have white dots at 7 o'clock_x000D_ 2. FCT was restricted_x000D_ 3. FSA was damaged_x000D_ 4. BSA was deformed |  | 1.Leaked at RC switch_x000D_ 2.Bending Up harden_x000D_ 3.FCT blocked_x000D_ _x000D_
|  |  | Parts | INSERTING SECTION SLEEVE ROHS NUT ASSY | FITTING PART FCT-G391A CHA-G391A AIR/WATER TUBE ASS''Y J TUBE LG-G391A FSA-G391A ROHS BSA G391A NUT ASSY SLEEVE ROHS | BSA G391A RBS-G181A FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE SLEEVE ROHS NUT ASSY RC SWITCH
|  |  | Approval | Cancel | Approval | Approval
46 | EG-600WR | 2G391K971 | Receive | 2018-09-06 10:22 | 2019-08-01 10:14 | 2019-10-17 13:39 | 2021-03-10 10:20 | 2021-11-12 11:09 | 2022-03-16 17:49
|  |  | Repair | 2019-03-05 16:56 | 2019-08-09 11:44 | 2019-10-17 14:57 | 2021-03-11 10:47 |  | 2022-03-21 11:31
|  |  | Size | Major | Minor | Minor | Minor | Major | Major
|  |  | Issue | FCT, ISA, LGB | OSA | VCA | BSA | FSA, BSA | FSA, BSA
|  |  | RMA# | FMSV2018090008 | FMSV2019080001 | FMSV2019100023 | FMSV2021030037 | FMSV2021110043 | FMSV2022030030
|  |  | Defect | FCT clogged at BSA portion_x000D_ CHA damaged by fluid_x000D_ WGB have fluid inside_x000D_ BSA wires loosen_x000D_ | UP wire is loosing | U/D was loosing_x000D_ Slide sleeve worn out | 1. FSA damaged at 40cm_x000D_ 2. BSA deformed_x000D_ 3. BSA loosen_x000D_ => Let customer temporary use | 1. FSA damaged_x000D_ 2. BSA deformed_x000D_ 3. fct restricted | 1. FSA damaged_x000D_ 2. BSA deformed_x000D_ 3. fct restricted
|  |  | Parts | SLEEVE ROHS NUT ASSY ROHS SLIDE SLEEVE CHA-FG391A ROHS TUBE(WJ) FG391A ROHS AW-TUBE FG391A ROHS RUBBER FOR BSA ROHS BENDING SECTION ASSY FCT RING FG381A-6 ROHS FCT ASSY FG391A ROHS |  | SLIDE SLEEVE |  | FSA-G391A ROHS ISA RING 2 ISA RING 1 BSA G391A RBS-G181A SLEEVE ROHS NUT ASSY FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE | FSA-G391A ROHS ISA RING 2 ISA RING 1 BSA G391A RBS-G181A SLEEVE ROHS NUT ASSY FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE
|  |  | Approval | Approval | Approval | Approval | Approval | Decline | Approval
47 | EG-600WR | 2G391K975 | Receive | 2019-10-08 11:10 | 2020-04-08 08:59 | 2020-04-28 15:19 | 2020-06-30 12:45 | 2020-09-08 13:18 | 2021-04-14 10:12
|  |  | Repair |  | 2020-04-09 16:35 |  | 2020-07-01 15:27 | 2020-09-16 08:44
|  |  | Size | Major | Major | Other | Minor | Major | Major
|  |  | Issue | CHA | CHA | O | OSA | DRY, BSA | CHA, FSA
|  |  | RMA# | FMSV2019100009 | FMSV2020040003 | FMSV2020040028 | FMSV2020060074 | FMSV2020090015 | FMSV2021040039
|  |  | Defect |  |  |  |  | 1. Lack glue on lens of CHA._x000D_ 2. BSA defromed. | 1. Image failure due to objective lens cracked_x000D_ 2. FSA scratched
|  |  | Parts | CAMERA HEAD ASS'Y TUBE(WJ) FG391A ROHS AW-TUBE FG391A ROHS FCT ASSY FG391A ROHS FITTING PART BENDING SECTION ASSY RBS-G181A SLEEVE ROHS NUT ASSY ROHS | CAMERA HEAD ASS'Y TUBE(WJ) FG391A ROHS AW-TUBE FG391A ROHS FCT ASSY FG391A ROHS FITTING PART BENDING SECTION ASSY RBS-G181A SLEEVE ROHS NUT ASSY |  |  |  | CHA-G391A AIR/WATER TUBE ASS''Y J TUBE FSA-G391A ROHS ISA RING 2 ISA RING 1 RBS-G181A CAP
|  |  | Approval | Cancel | Approval | Cancel | Approval | Approval | Cancel
48 | EG-600WR | 2G391K977 | Receive | 2019-11-04 11:47 | 2021-06-10 14:09 | 2021-07-26 15:14 | 2021-08-30 12:49 | 2022-02-23 12:16 | 2022-05-05 12:22
|  |  | Repair |  | 2021-06-16 07:56 | 2021-08-12 09:31 |  | 2022-02-24 17:39
|  |  | Size | Major | Major | Minor | Minor | Minor | Minor
|  |  | Issue | CHA, FCT, BSA | FCT | O, WGB, SW | DRY, SW | SW | O
|  |  | RMA# | FMSV2019110005 | FMSV2021060004 | FMSV2021070038 | FMSV2021080005 | FMSV2022020028 | FMSV2022050011
|  |  | Defect | 1. Black spots_x000D_ 2. FCT leaking_x000D_ 3. BSA deformed and loosen |  | 1. Stain image_x000D_ 2. LG lens glue damaged_x000D_ 3. FR switch broken. | 1.Image stain._x000D_ 2.Black spots at 7o'clock._x000D_ 3.Leaked at RC switch.
|  |  | Parts | INSERTING SECTION SLEEVE ROHS NUT ASSY ROHS | FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE RBS-G181A | FR SWITCH | RC SWITCH FR SWITCH RBS-G181A RUBBER CONE SET | RC SWITCH FR SWITCH RBS-G181A
|  |  | Approval | Cancel | Approval | Approval | Cancel | Approval | NoFault
49 | EG-600WR | 2G391K978 | Receive | 2020-01-10 13:39 | 2020-02-03 09:51 | 2021-04-26 13:55 | 2021-05-26 15:41 | 2022-03-01 18:12
|  |  | Repair | 2020-01-15 15:55 | 2020-04-17 17:15 | 2021-05-14 10:31 | 2021-06-25 12:17 | 2022-03-18 16:10
|  |  | Size | Minor | Major | Major | Minor | Minor
|  |  | Issue | BSA | FCT | AWT, DRY, FSA | OSA | RBS, OSA, SUV, VCA
|  |  | RMA# | FMSV2020010013 | FMSV2020020003 | FMSV2021040079 | FMSV2021050034 | FMSV2022030001
Defect
|  |  | Parts |  | FCT ASSY FG391A ROHS FITTING PART RBS-G181A | RBS-G181A FSA-G391A ROHS ISA RING 2 ISA RING 1 |  | SLIDE SLEEVE VALVE RING(RED) FITTING PART ASSY ROHS RBS-G181A
|  |  | Approval | Approval | Approval | Approval | Approval | Approval
50 | EG-600WR | 2G391K979 | Receive | 2021-03-29 13:22 | 2021-07-01 14:02 | 2022-04-25 13:56
|  |  | Repair | 2021-03-31 09:32 | 2021-07-29 17:36
|  |  | Size | Minor | Major | Minor
|  |  | Issue | ANGL, BSA | FCT, BSA, WGB | O
|  |  | RMA# | FMSV2021030106 | FMSV2021070003 | FMSV2022040063
|  |  | Defect | UD LOOSEN, BSA deformation | 1.Leaked at FCT._x000D_ 2.Bending deformation._x000D_ 3.Missing glue around LG lens.
|  |  | Parts |  | BSA G391A RBS-G181A FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE SLEEVE ROHS NUT ASSY SLIDE SLEEVE
|  |  | Approval | Approval | Approval | NoFault
51 | EG-600WR | 2G391K980 | Receive | 2019-11-04 11:48 | 2021-06-10 17:56 | 2022-03-23 13:04
|  |  | Repair |  | 2021-06-16 11:36 | 2022-03-25 18:49
|  |  | Size | Major | Major | Minor
|  |  | Issue | CHA, BSA, LGB, FCT, O, VCA | DRY, BSA, LGB, FCT, O, VCA | ANGL
|  |  | RMA# | FMSV2019110006 | FMSV2021060009 | FMSV2022030043
|  |  | Defect | White spots appears at 7 o'clock on the image._x000D_ Up wire of BSA was broken._x000D_ FCT is restricted ar BSA._x000D_ Missing glue around OBL_x000D_ LGB has been broken up to 25%. _x000D_ | White spots appears at 7 o'clock on the image._x000D_ Up wire of BSA was broken._x000D_ FCT is restricted ar BSA.
|  |  | Parts | INSERTING SECTION SLEEVE ROHS NUT ASSY ROHS SLIDE SLEEVE | FCT-G391A FITTING PART AIR/WATER TUBE ASS''Y J TUBE LG-G391A BSA G391A RBS-G181A SLEEVE ROHS NUT ASSY SLIDE SLEEVE
|  |  | Approval | Cancel | Approval | Approval
52 | EG-600WR | 2G391K982 | Receive | 2020-08-05 13:58 | 2021-04-07 10:38 | 2021-11-12 11:06
|  |  | Repair |  | 2021-04-12 08:34 | 2021-11-17 15:41
|  |  | Size | Major | Major | Minor
|  |  | Issue | BSA | BSA, FCT, AWT, WJT, DEC | OSA, WGA/WGB
|  |  | RMA# | FMSV2020080009 | FMSV2021040008 | FMSV2021110042
|  |  | Defect |  | BSA damaged_x000D_ FSA damaged_x000D_ CAP DAMAGED | 1. U/D loosen_x000D_ 2. WGB glue damaged_x000D_ Note: BSA deformed
|  |  | Parts | BSA G391A RBS-G181A FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE SLEEVE ROHS NUT ASSY COVER PLATE | FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE BSA G391A RBS-G181A COVER PLATE LR LOCK KNOB E-CLIP NUT ASSY SLEEVE ROHS
|  |  | Approval | Decline | Approval | Approval
53 | EG-600WR | 2G391K985 | Receive | 2020-04-28 08:59 | 2020-09-29 08:15 | 2020-11-02 14:17 | 2020-12-18 15:48 | 2021-03-11 10:32
|  |  | Repair |  |  | 2020-12-14 16:43
|  |  | Size | Major | Maintain | Major |  | Major
|  |  | Issue | CHA, FCT, FSA, RBS, WGB, VCA | O | CHA, FCT, FSA, RBS, BSA | O | CHA, FCT, VCA
|  |  | RMA# | FMSV2020040027 | FMSV2020090061 | FMSV2020110006 | FMSV2020120045 | FMSV2021030045
|  |  | Defect | 1) Black spots appears on images at 7 o'clock. _x000D_ 2) FCT restricted at bending section. _x000D_ 3) FSA surface is peeling _x000D_ 4) Missing glue around WGB. _x000D_ 5. RBS glue cracked _x000D_ 6) Belt of VCA broken _x000D_ 7) Air leak at RBS. |  | 1) Black spots, stain appear on images at 1 o'clock. _x000D_ 2) FCT restricted at the bending section. _x000D_ 3) FSA surface is peeling _x000D_ 4) Missing glue around WGB. _x000D_ 5. RBS glue cracked _x000D_ 6) Belt of VCA broken _x000D_ 7) Air leak at RBS_x000D_ 8) BSA broken | Scope use normally | FCT Restriction, cracked on camera glass
|  |  | Parts | CAMERA HEAD ASS'Y FCT-G391A FITTING PART AIR/WATER TUBE ASS''Y J TUBE CAP FSA-G391A ROHS RING FG378A-6 ROHS FSA RING BELT RBS-G181A |  | BSA G391A SLEEVE ROHS NUT ASSY CAMERA HEAD ASS'Y FCT-G391A FITTING PART AIR/WATER TUBE ASS''Y J TUBE CAP FSA-G391A ROHS RING FG378A-6 ROHS FSA RING BELT RBS-G181A |  | CAMERA HEAD ASS'Y J TUBE AIR/WATER TUBE ASS''Y FCT-G391A FITTING PART RBS-G181A
|  |  | Approval | Cancel | Cancel | Approval | Cancel | Cancel
54 | EG-600WR | 2G391K986 | Receive | 2020-09-16 10:25 | 2020-10-29 14:21 | 2021-06-15 15:16 | 2021-09-13 17:31
|  |  | Repair | 2020-10-02 09:05 | 2020-11-03 10:14 | 2021-07-01 13:43
|  |  | Size | Major | Maintain | Major | Major
|  |  | Issue | FSA | BSA | LGB, BSA, FSA, FCT | FSA, BSA, FCT, VCA, LGB
|  |  | RMA# | FMSV2020090037 | FMSV2020100067 | FMSV2021060035 | FMSV2021090019
|  |  | Defect | FCT LEAKED |  |  | 1. FSA scratched_x000D_ 2. BSA deformed_x000D_ 3. FCT scratched_x000D_ 4. Belt of VCA cap broken_x000D_ 5. LGB broken 30% (1220 lx)_x000D_ 6. LG lens glue damaged_x000D_ Note: No leakage
|  |  | Parts | FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE RBS-G181A |  | FITTING PART FCT-G391A AIR/WATER TUBE ASS''Y J TUBE LG-G391A SLEEVE ROHS NUT ASSY BSA G391A RBS-G181A RING2 FG391A ROHS FSA-G391A ROHS ISA RING 2 ISA RING 1 | FCT-G391A FITTING PART AIR/WATER TUBE ASS''Y J TUBE LG-G391A BSA G391A RBS-G181A FSA-G391A ROHS RING2 FG391A ROHS ISA RING 2 O-RING ISA RING 1 SLEEVE ROHS NUT ASSY
|  |  | Approval | Approval | Approval | Approval | Decline
55 | VP-7000 | 2V644K158 | Receive | 2021-12-22 11:47
|  |  | Repair | 2021-12-22 13:07
|  |  | Size | Minor
|  |  | Issue | SOF
|  |  | RMA# | FMSV2021120051
Defect
Parts
|  |  | Approval | Approval
56 | ED-530XT8 | 3D103L050 | Receive | 2019-11-06 15:16 | 2020-05-19 10:33 | 2020-08-24 15:07 | 2023-05-31 10:30 | 2023-06-27 10:56 | 2023-11-02 15:48
|  |  | Repair |  | 2020-05-19 12:58 |  | 2023-06-07 15:57 | 2023-10-12 12:24
|  |  | Size | Maintain | Minor | Maintain | Major | Major | Major
|  |  | Issue | O | OSA | O | FEW | FCT | FCT
|  |  | RMA# | FMSV2019110013 | FMSV2020050041 | FMSV2020080048 | FMSV2023050080 | FMSV2023060027 | FMSV2023110004
|  |  | Defect |  | UD, LR loosen | Scope OK |  | FCT leaked | 1. FCT leaked
|  |  | Parts |  |  |  | ELEVATOR WIRE ASS/Y ROHS SPECIAL SCREW SLEEVE | FITTING PART ROHS FCT-D103A RUBBER FOR BSA ROHS A/W TUBE ASS''Y | FITTING PART ROHS FCT-D103A RUBBER FOR BSA ROHS
|  |  | Approval | NoFault | Approval | Inspection | Approval | Approval
57 | ED-530XT8 | 3D103L052 | Receive | 2019-11-06 16:31 | 2020-08-25 09:32 | 2023-05-31 10:41 | 2023-08-10 12:29
|  |  | Repair | 2020-01-10 10:44 | 2020-08-25 11:54 |  | 2023-08-17 16:29
|  |  | Size | Major | Minor | Maintain | Major
|  |  | Issue | CHA, FCT | OSA | O | CHA
|  |  | RMA# | FMSV2019110015 | FMSV2020080051 | FMSV2023050044 | FMSV2023080037
|  |  | Defect | FCT restriction at 5cm_x000D_ CHA has black spot | BSA wires LOOSEN | QC | Black spots
|  |  | Parts | CHA-D102A FITTING PART ROHS FCT-D103A RUBBER FOR BSA ROHS |  |  | CHA-D102A RUBBER FOR BSA ROHS
|  |  | Approval | Approval | Approval | Inspection | Approval
58 | ED-530XT8 | 3D103L053 | Receive | 2021-07-08 11:09 | 2022-03-16 12:44 | 2022-07-27 17:38 | 2023-09-08 13:50
|  |  | Repair |  | 2022-03-18 13:05 | 2022-08-01 16:37 | 2023-09-14 19:07
|  |  | Size | Major | Minor | Minor | Minor
|  |  | Issue | DTA, CHA, FSA, BSA | DTA, O | OSA | CLN, DRY
|  |  | RMA# | FMSV2021070017 | FMSV2022030025 | FMSV2022070102 | FMSV2023090010
|  |  | Defect | 1. DTA leaked_x000D_ 2. Image stain_x000D_ 3. FSA deformed_x000D_ 4. BSA stiff | - Stain on image_x000D_ - Cover DTA leaked | U/D loosen
|  |  | Parts | ISA-D103A PIPE SLEEVE NUT ASSY SLEEVE SPECIAL SCREW
|  |  | Approval | Decline | Approval | Approval | Approval
59 | ED-530XT8 | 3D103L054 | Receive | 2020-07-02 11:12 | 2021-10-11 14:00
Repair
|  |  | Size | Maintain | Major
|  |  | Issue | O | FSA, ISA
|  |  | RMA# | FMSV2020070011 | FMSV2021100018
|  |  | Defect |  | FSA leaked_x000D_ Image noise_x000D_ BSA loosen
|  |  | Parts |  | ISA-D103A PIPE SLEEVE NUT ASSY
|  |  | Approval | Inspection | Cancel
60 | XL-4450 | 3S094D287 | Receive | 2022-04-15 09:56
|  |  | Repair | 2022-04-21 15:15
|  |  | Size | Maintain
|  |  | Issue | O
|  |  | RMA# | FMSV2022040028
|  |  | Defect | XL4450 is working normally
Parts
|  |  | Approval | Approval
61 | XL-4450 | 3S094D288 | Receive | 2021-05-04 15:47
|  |  | Repair | 2021-05-26 09:32
|  |  | Size | Other
|  |  | Issue | LMP, O
|  |  | RMA# | FMSV2021050006
|  |  | Defect | dark lights
|  |  | Parts | LMP-402
|  |  | Approval | Approval
62 | BL-7000 | 3S101K027 | Receive | 2021-01-18 15:58
|  |  | Repair | 2021-01-22 08:29
|  |  | Size | Other
|  |  | Issue | PUM, O
|  |  | RMA# | FMSV2021010029
|  |  | Defect | Air pump weak
|  |  | Parts | Pump Assy
|  |  | Approval | Approval
63 | VP-4450HD | 3V567D283 | Receive | 2021-03-22 16:10 | 2022-05-10 17:37
|  |  | Repair | 2021-03-25 11:37 | 2022-05-18 15:13
|  |  | Size | Other | Maintain
|  |  | Issue | O | O
|  |  | RMA# | FMSV2021030079 | FMSV2022050020
|  |  | Defect |  | VP4450 is working normally
Parts
|  |  | Approval | Approval | Approval
64 | VP-4450HD | 3V567D284 | Receive | 2021-09-23 19:29
|  |  | Repair | 2021-10-28 15:39
|  |  | Size | Other
|  |  | Issue | O
|  |  | RMA# | FMSV2021090035
Defect
Parts
|  |  | Approval | Approval
65 | EG-530UR2 | RU039A099 | Receive | 2019-08-09 14:46
Repair
|  |  | Size | Major
|  |  | Issue | ISA
|  |  | RMA# | FMSV2019080047
Defect
|  |  | Parts | ISA-U39A SLEEVE ROHS NUT ASSY ROHS
|  |  | Approval | Decline
66 | PB-20 | RW512A161 | Receive | 2022-03-17 17:21
Repair
|  |  | Size | Other
|  |  | Issue | O
|  |  | RMA# | FMSV2022030031
|  |  | Defect | This PB20 is working normally.
Parts
|  |  | Approval | Inspection

## Sheet1
RMA No. | SERIAL_NO | MODEL | RECIEVE_DATE | Repair Date | Repair Size | Defect Note | Approval | Repair | No#
FMSV2023120017 | 1C610D007 | EC-590ZW3/M | 2023-12-12 18:58:48 |  | Major | FSB damaged FCT leaked BSA loosen | Approval | FSB, FCT, OSA | 1 | Broken
FMSV2019080003 | 1C657K030 | EC-530LP | 2019-08-01 10:57:15 | 2019-08-07 08:25:16 | Minor | VCA CALBLE BROKEN | Approval | VCA | 1 | OK
FMSV2023040004 | 1C675K021 | EN-580T | 2023-04-03 17:13:24 |  | Major | 1.Leaked at FCT 2.BSA Stiff 3.Damaged glue at DTE | Cancel | FCT, BSA | 4 | Broken
FMSV2020120080 | 1C675K022 | EN-580T | 2020-12-30 10:48:17 |  | Major | 1. Bending Left broken. 2. Image black spot at 1 o'clock. 3. FCT leak and block. | Cancel | DRY, BSA, FCT | 2 | Broken
FMSV2018050055 | 1C692K117 | EC-600WI | 2019-02-26 16:18:27 |  | Major | Blue image, black spot at 7 o'clock FCT was restricted | Cancel | CHA, FCT | 1 | Broken
FMSV2019020016 | 1C692K118 | EC-600WI | 2019-02-28 09:50:25 |  | Major | Black spot on image Air water jet clogged Wires lossen LGB broken about 20% | Cancel | CHA, LGB, BSA, AWT | 2 | Broken
FMSV2023070076 | 1C692K141 | EC-600WL | 2023-07-31 12:31:16 |  | Major | 1.Leaked at FSB 2.VCA connector worn out | Cancel | FSB, VCA | 6 | Broken
FMSV2019100055 | 1C692K142 | EC-600WI | 2019-10-28 16:08:48 |  | Major | BSA deformed AWT clogged | Cancel | BSA, AWT | 3 | Broken
FMSV2020110016 | 1C692K159 | EC-600WL | 2020-11-03 16:27:57 |  | Major | 1. Image failed at 5 o'clock. 2. FSA worn out. 3. VCA connector leaked and loosen. | Cancel | CHA, FSA, VCA | 2 | Broken
FMSV2021040040 | 1C692K651 | EC-600WI | 2021-04-14 11:32:42 |  | Maintain | not received | Decline | E | 6 | Broken
FMSV2022040027 | 1C692K802 | EC-600WI | 2022-04-14 18:00:17 |  | Major | FSB broken BSA wire loosen. | Transfer | FSB | 6 | Broken
FMSV2021100043 | 1C692K803 | EC-600WI | 2021-10-25 18:06:02 | 2021-12-09 17:06:05 | Major | Blue Image | Approval | CHA | 7 | OK
FMSV2020110033 | 1C692K806 | EC-600WI | 2020-11-03 17:01:00 |  | Major | 1. FCT closed 2. BSA deformed | Cancel | FSA, BSA | 4 | Broken
FMSV2022040023 | 1C692K826 | EC-600WI | 2022-04-13 18:58:09 |  | Maintain | Inspection OK | Transfer | O | 5 | Broken
FMSV2022040030 | 1C692K833 | EC-600WI | 2022-04-15 10:42:52 | 2022-04-19 15:39:58 | Minor | Image noise due to VCA cable broken | Approval | VCA | 7 | OK
FMSV2021060011 | 1C692K835 | EC-600WI | 2021-06-10 17:57:28 |  | Major | 1. FSA leaked 2. FCT restricted 3. BSA stiff | Decline | FSA, FCT, BSA | 7 | Broken
FMSV2020090048 | 1C692K836 | EC-600WI | 2020-09-21 10:32:11 | 2020-09-23 08:11:44 | Minor |  | Approval | ANGL | 2 | OK
FMSV2019090030 | 1C730K096 | EC-760ZP-V/M | 2019-09-16 17:08:52 |  | Major |  | Decline | ANGL, CHA | 1 | Broken
FMSV2019100008 | 1G349K025 | EG-590ZW | 2019-10-08 11:08:44 |  | Major | blur and stain image, FCT restricted, VCA scratch, LGB broken 40% | Cancel | CHA, FCT, VCA, LGB | 1 | Broken
FMSV2019120009 | 1G391K219 | EG-600WR | 2019-12-05 12:04:34 |  | Major |  | Cancel | CHA, BSA | 3 | Broken
FMSV2019080045 | 1G391K223 | EG-600WR | 2019-08-09 14:41:30 |  |  | CHA len broken, Fct restriction, bsa deformed. VCA leaking | Decline | CHA, FCT, VCA, BSA | 1 | Broken
FMSV2019080044 | 1G391K324 | EG-600WR | 2019-08-09 14:39:55 |  | Major | 1. Stain image 2. FCT restriction 3. BSA deformed 4. VCA loosen | Decline | CHA, BSA, FCT, VCA | 1 | Broken
FMSV2019050001 | 1G391K366 | EG-600WR | 2019-05-06 15:10:27 |  | Major | Up wire broken VCA damaged | Cancel | VCA, BSA | 6 | Broken
FMSV2019080046 | 1G391K368 | EG-600WR | 2019-08-09 14:43:25 |  | Major | 1. White spots and stain image at 7 o'clock. 2. FCT is restriction at bending section. 3. FSA is deformed. 4. Leaking at MM Switch. 5. BSA is deformed. | Decline | CHA, FCT, FSA, SW, BSA | 1 | Broken
FMSV2018110059 | 1G391K420 | EG-600WR | 2018-11-26 15:30:23 |  | Major |  | Cancel | LPA, BSA, VCA | 5 | Broken
FMSV2022010037 | 1G402K918 | EG-760R | 2022-01-25 17:08:01 | 2022-02-17 15:06:18 | Major | leaked test, CHA black spot | Approval | FCT, DRY, FSA | 5 | OK
FMSV2021010033 | 1G403K346 | EG-760Z | 2021-01-18 09:30:23 |  | Major | noise on image scope leak | Decline | AWT, ISA, LC | 3 | Broken
FMSV2023080081 | 1S094K191 | XL-4450 | 2023-08-21 17:26:30 |  | Other |  | Cancel | IRS | 2 | Broken
FMSV2023020009 | 1S094K192 | XL-4450 | 2023-02-02 18:13:18 | 2023-04-12 10:27:12 | Other | Iris failure | Approval | IRS | 5 | OK
FMSV2022040042 | 1S094K193 | XL-4450 | 2022-04-20 13:56:44 | 2022-04-20 19:41:44 | Other | 1. Movements failure 2. Wear air | Approval | IRS, O | 3 | OK
FMSV2022110016 | 1S094K194 | XL-4450 | 2022-11-08 18:45:45 | 2022-11-10 19:32:28 | Minor | Air weak | Approval | PUM | 3 | OK
FMSV2023020007 | 1S094K195 | XL-4450 | 2023-02-02 18:11:48 | 2023-04-12 10:27:29 | Other | Iris failure | Approval | IRS | 4 | OK
FMSV2022020014 | 1V567K309 | VP-4450HD | 2022-02-17 11:15:25 |  | Other |  | Cancel | PCB | 1 | Broken
FMSV2023010004 | 1V567K312 | VP-4450HD | 2023-01-03 13:48:28 |  | Other |  | Decline | PCB | 1 | Broken
FMSV2021030046 | 2C667D037 | EC-590ZW3/L | 2021-03-11 10:34:39 | 2021-03-18 17:11:29 | Minor | 1. bsa loosen 2. awj restricted 3. Glue of camera len broken | Approval | BSA, AWT | 1 | OK
FMSV2022020030 | 2C728K101 | EC-760R-V/I | 2022-02-24 13:00:15 | 2022-02-25 13:15:59 | Minor |  | Approval | LC | 5 | OK
FMSV2023090057 | 2C728K102 | EC-760R-V/I | 2023-09-25 18:34:50 | 2023-09-28 12:08:32 | Minor |  | Approval | BSA | 8 | OK
FMSV2023020022 | 2C730K032 | EC-760ZP-V/M | 2023-02-08 18:54:02 |  | Major | 1.LGB broken 2.BSA deformed 3.FSB worn out 4.Damaged glue at DEC | Cancel | LGB, BSA, FSB | 6 | Broken
FMSV2023090015 | 2G391K487 | EG-600WR | 2023-09-08 15:01:08 |  | Major | BSA heavy torque | Approval | BSA | 6 | Broken
FMSV2023120050 | 2G391K630 | EG-600WR | 2023-12-28 10:29:50 |  | Major | 1. FSB broken => leaked 2. Noise Image => drying => CHA OK 3. PCB damaged | Transfer | FSA, DRY, PCB | 10 | Broken
FMSV2021100039 | 2G391K955 | EG-600WR | 2021-10-22 15:12:13 |  | Major | FSA crrosion WJT leaked CHA has black spots | Cancel | FSA, WJT, DRY, CHA | 5 | Broken
FMSV2021050014 | 2G391K958 | EG-600WR | 2021-05-07 15:59:50 | 2021-05-13 18:02:39 | Minor |  | Approval | BSA | 6 | OK
FMSV2022040025 | 2G391K960 | EG-600WR | 2022-04-14 10:46:13 | 2022-04-18 12:04:36 | Major | RBS leak CHA black spots BSA worn out | Approval | RBS, DRY, CHA, BSA | 5 | OK
FMSV2022030032 | 2G391K963 | EG-600WR | 2022-03-17 18:00:05 | 2022-03-23 13:36:23 | Major | 1. BSA broken 2. CHA lens scratched 3. FSB broken 4. Belt of VCA broken 5. Noise image due to VCA cable broken | Approval | BSA, CLN, FSB, VCA | 6 | OK
FMSV2021100030 | 2G391K967 | EG-600WR | 2021-10-15 15:30:06 | 2022-02-11 11:06:05 | Major | 1.Leaked at RC switch 2.Bending Up harden 3.FCT blocked | Approval | SW, BSA, FCT | 3 | OK
FMSV2022030030 | 2G391K971 | EG-600WR | 2022-03-16 17:48:39 | 2022-03-21 11:30:51 | Major | 1. FSA damaged 2. BSA deformed 3. fct restricted | Approval | FSA, BSA | 6 | OK
FMSV2021040039 | 2G391K975 | EG-600WR | 2021-04-14 10:12:11 |  | Major | 1. Image failure due to objective lens cracked 2. FSA scratched | Cancel | CHA, FSA | 6 | Broken
FMSV2022050011 | 2G391K977 | EG-600WR | 2022-05-05 12:22:15 |  | Minor |  | NoFault | O | 6 | Broken
FMSV2022030001 | 2G391K978 | EG-600WR | 2022-03-01 18:12:10 | 2022-03-18 16:09:41 | Minor |  | Approval | RBS, OSA, SUV, VCA | 5 | OK
FMSV2022040063 | 2G391K979 | EG-600WR | 2022-04-25 13:55:44 |  | Minor |  | NoFault | O | 3 | Broken
FMSV2022030043 | 2G391K980 | EG-600WR | 2022-03-23 13:04:17 | 2022-03-25 18:48:44 | Minor |  | Approval | ANGL | 3 | OK
FMSV2021110042 | 2G391K982 | EG-600WR | 2021-11-12 11:06:11 | 2021-11-17 15:40:52 | Minor | 1. U/D loosen 2. WGB glue damaged Note: BSA deformed | Approval | OSA, WGA/WGB | 3 | OK
FMSV2021030045 | 2G391K985 | EG-600WR | 2021-03-11 10:32:09 |  | Major | FCT Restriction, cracked on camera glass | Cancel | CHA, FCT, VCA | 5 | Broken
FMSV2021090019 | 2G391K986 | EG-600WR | 2021-09-13 17:30:49 |  | Major | 1. FSA scratched 2. BSA deformed 3. FCT scratched 4. Belt of VCA cap broken 5. LGB broken 30% (1220 lx) 6. LG lens glue damaged Note: No leakage | Decline | FSA, BSA, FCT, VCA, LGB | 4 | Broken
FMSV2021120051 | 2V644K158 | VP-7000 | 2021-12-22 11:47:26 | 2021-12-22 13:06:45 | Minor |  | Approval | SOF | 1 | OK
FMSV2023110004 | 3D103L050 | ED-530XT8 | 2023-11-02 15:47:41 |  | Major | 1. FCT leaked |  | FCT | 6 | Broken
FMSV2023080037 | 3D103L052 | ED-530XT8 | 2023-08-10 12:29:00 | 2023-08-17 16:29:03 | Major | Black spots | Approval | CHA | 4 | OK
FMSV2023090010 | 3D103L053 | ED-530XT8 | 2023-09-08 13:49:57 | 2023-09-14 19:07:11 | Minor |  | Approval | CLN, DRY | 4 | OK
FMSV2021100018 | 3D103L054 | ED-530XT8 | 2021-10-11 13:59:48 |  | Major | FSA leaked Image noise BSA loosen | Cancel | FSA, ISA | 2 | Broken
FMSV2022040028 | 3S094D287 | XL-4450 | 2022-04-15 09:56:03 | 2022-04-21 15:14:32 | Maintain | XL4450 is working normally | Approval | O | 1 | OK
FMSV2021050006 | 3S094D288 | XL-4450 | 2021-05-04 15:47:26 | 2021-05-26 09:31:46 | Other | dark lights | Approval | LMP, O | 1 | OK
FMSV2021010029 | 3S101K027 | BL-7000 | 2021-01-18 15:58:29 | 2021-01-22 08:29:25 | Other | Air pump weak | Approval | PUM, O | 1 | OK
FMSV2022050020 | 3V567D283 | VP-4450HD | 2022-05-10 17:36:49 | 2022-05-18 15:12:47 | Maintain | VP4450 is working normally | Approval | O | 2 | OK
FMSV2021090035 | 3V567D284 | VP-4450HD | 2021-09-23 19:28:33 | 2021-10-28 15:38:54 | Other |  | Approval | O | 1 | OK
FMSV2019080047 | RU039A099 | EG-530UR2 | 2019-08-09 14:45:38 |  | Major |  | Decline | ISA | 1 | Broken
FMSV2022030031 | RW512A161 | PB-20 | 2022-03-17 17:20:59 |  | Other | This PB20 is working normally. | Inspection | O | 1 | Broken